## Maintenance Log
NOTICE: This document is an example only. It must be revised to reflect your lab’s specific processes and/or specific protocol requirements.
Approved and current. Effective starting 07-Feb-2024. Last reviewed on 07-Feb-2024. RDP 204 (version 1.0). Analytical Balance Daily/Annual Maintenance and Calibration Logs. Page 1 of 6
| Laboratory Name | Department Name
Analytical Balance Maintenance Log
| Manufacturer: | Balance Location:
| Balance Model: | Date:
Balance Serial Number:
Date
Start Up
Clean the balance
Auto Calibration
Daily Calibration
Shut Down
Tech Initials
Date
Start Up
Clean the balance
Auto Calibration
Daily Calibration
Shut Down
Tech Intials
Date
Start Up
Clean the balance
Auto Calibration
Daily Calibration
Shut Down
Tech Intials
| Reviewed By ______________________________ |  | Date __________________

## Daily Calibration Log
NOTICE: This document is an example only. It must be revised to reflect your lab’s specific processes and/or specific protocol requirements.
Approved and current. Effective starting 07-Feb-2024. Last reviewed on 07-Feb-2024. RDP 204 (version 1.0). Analytical Balance Daily/Annual Maintenance and Calibration Logs. Page 2 of 6
| Laboratory Name |  |  | Laboratory Section
Analytical Balance Calibration Check Log
| Manufacturer: |  |  | Balance Location:
Balance Model:
| Balance Serial Number: |  |  | Date:
Standard Weight: _______________________________
Acceptance Limits: _______________________________
| Date | Weight | Acceptable/ Not Acceptable | Comments | Tech Initials
| Reviewed by:_________________________ |  |  | Date:____________________

## Reproducibility Log
NOTICE: This document is an example only. It must be revised to reflect your lab’s specific processes and/or specific protocol requirements.
Approved and current. Effective starting 07-Feb-2024. Last reviewed on 07-Feb-2024. RDP 204 (version 1.0). Analytical Balance Daily/Annual Maintenance and Calibration Logs. Page 3 of 6
| Laboratory Name |  | Laboratory Section
Annual Reproducibility Check
| Manufacturer: |  | Balance Location:
| Balance Model: |  | Reference Weight Serial Number:
| Balance Serial Number: |  | Date:
| Measurement | Mass of Reference Weight, g | Mass When Reference Weight is Removed, g | Delta (Reference Weight, g-Reference Weight Removed, g)
| 1 |  |  | 0
| 2 |  |  | 0
| 3 |  |  | 0
| 4 |  |  | 0
| 5 |  |  | 0
| 6 |  |  | 0
| 7 |  |  | 0
| 8 |  |  | 0
| 9 |  |  | 0
| 10 |  |  | 0
| Mean (average) delta: |  |  | 0
| SD of the 10 delta values: |  |  | 0
Manufacturer's specification (SD):
| SD is within acceptable tolerance range: |  |  | YES/NO
| Performed By: |  |  | Date:
| Reviewed By: |  |  | Date:

## Annual Calibration Check
NOTICE: This document is an example only. It must be revised to reflect your lab’s specific processes and/or specific protocol requirements.
Approved and current. Effective starting 07-Feb-2024. Last reviewed on 07-Feb-2024. RDP 204 (version 1.0). Analytical Balance Daily/Annual Maintenance and Calibration Logs. Page 4 of 6
|  | Laboratory Name |  | Laboratory Section
Annual Calibration Check
|  | Manufacturer: |  | Balance Location:
Balance Model:
|  | Balance Serial Number: |  | Date:
Date _______________
Date of the Previous Calibration Check ______________
|  | Test Weight | Weight at previous Calibration | Weight at Current Calibration | Acceptable/ unacceptable | Tech Initials
1
2
3
4
5
6
Reviewed by:___________________________ Date: ________________
Date _______________
Date of the Previous Calibration Check ______________
|  | Test Weight | Weight at previous Calibration | Weight at Current Calibration | Acceptable/ unacceptable | Tech Initials
1
2
3
4
5
6
Reviewed by:___________________________ Date: ________________

## Linearity&Cornerload
NOTICE: This document is an example only. It must be revised to reflect your lab’s specific processes and/or specific protocol requirements.
Approved and current. Effective starting 07-Feb-2024. Last reviewed on 07-Feb-2024. RDP 204 (version 1.0). Analytical Balance Daily/Annual Maintenance and Calibration Logs. Page 5 of 6
|  | Laboratory Name |  | Laboratory Section
Linearity Test
|  | Manufacturer: |  | Balance Location:
Balance Model:
|  | Balance Serial Number: |  | Date:
Date: ____________
Weight of Test Weight A : ___________________
Weight of Test Weight B: ____________________
|  | Weight: 0% - 50% | Weight: 50%-100% | Difference | Acceptable/ Unacceptable | Tech Initials
Reviewed by:_________________________ Date: ________________
Cornerload Test
Date: ____________
Test Weight: ___________________
Cornerload Check
|  |  | Weight | Acceptable/ Unacceptable | Tech Initials
Weight at Center of Pan
Weight at Front of Pan
Weight at Right of Pan
Weight at Left of Pan
Weight at Rear of Pan
Reviewed by:_________________________ Date: ________________

## Recommended Tolerance Limits
NOTICE: This document is an example only. It must be revised to reflect your lab’s specific processes and/or specific protocol requirements.
Approved and current. Effective starting 07-Feb-2024. Last reviewed on 07-Feb-2024. RDP 204 (version 1.0). Analytical Balance Daily/Annual Maintenance and Calibration Logs. Page 6 of 6
Recommended Tolerance Limits
| Laboratory Name |  |  | Laboratory Section
Recommended Acceptability Limits for Analytical Balances - use if instrument manufacturer ranges are not available
Laboratory Balance Calibration
Certified Weights >500g =+/- 0.5g
Certified Weights 100 - 500g = +/- 0.1g
Certified Weights <100g = +/- 0.01g
Laboratory Balance Reproducibility
Reproducibility (SD) = 0.1 mg
Laboratory Balance Linearity
Linearity = +/- 0.2 mg
Laboratory Cornerload Tolerances
Resolution
|  | 0.1g | .01g | .001g | .0001g | .00001g
Capacity
| 30g | * | * | * | 2 | 5
| 100g | * | * | 2 | 4 | 10
| 300g | * | 2 | 4 | 10 | *
| 1000g | 2 | 4 | 10 | * | *
The value for each combination of balance capacity and resolution is the greatest allowable difference between the highest and lowest readings of the cornerload test weights in the most right-hand digits of the balance readout.