{"text_format":"2","workbook":"xlsx","n_sheets":4,"sheets":["INTRAENSAYO CN-HBsAg II","INTRAENSAYO CP-HBsAg II","INTERENSAYO CN HBsAg","INTERENSAYO CP HBsAg"],"not_rendered":{"charts":0,"images":0,"macros":0,"external_links":0,"defined_names":0,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"run_cost":"one","rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"INTRAENSAYO CN-HBsAg II","index":0,"n_rows":46,"n_cols":5,"rows_kept":46,"cut":false,"hidden":false,"merged":[[37,3,38,3],[37,2,38,2],[37,4,38,4],[2,1,2,4],[3,1,3,4]]}
["NOTICE: This document is an example only. It must be revised to reflect your lab’s specific processes and/or specific protocol requirements."]
["Approved and current. Effective starting 10-Aug-2023. Last reviewed on 10-Aug-2023.\r\nRDP 508 (version 1.0). Precision (AngloLab HbsAg)_Example. Page 1 of 4"]
[null,"ANGLOLAB S.A"]
[null,"CLINICA ANGLO AMERICANA"]
["           PRUEBA DE PRECISION INTRAENSAYO COBAS PURE E402"]
[]
["PRUEBA: HBsAg II",null,"Fecha: 11 de enero 2023"]
[]
["CONTROL : PC HBsAg II 1",null,"Lote: 55190199",null,"Expira: 28-02-2023"]
["MARCA REACTIVO: ROCHE",null,"Lote kit: 64979002",null,"Expira: 31-05-2024"]
["MARCA CALIBRADOR: ROCHE",null,"Lote kit: 64979002",null,"Expira: 31-05-2024"]
[]
[null,null,"CORRIDA","PC HBsAg II 1      (COI)"]
[null,null,1,0.39200000000000002]
[null,null,2,0.40699999999999997]
[null,null,3,0.39800000000000002]
[null,null,4,0.371]
[null,null,5,0.38]
[null,null,6,0.47199999999999998]
[null,null,7,0.377]
[null,null,8,0.39400000000000002]
[null,null,9,0.39400000000000002]
[null,null,10,0.42799999999999999]
[null,null,11,0.41699999999999998]
[null,null,12,0.41099999999999998]
[null,null,13,0.39500000000000002]
[null,null,14,0.40500000000000003]
[null,null,15,0.377]
[null,null,16,0.38400000000000001]
[null,null,17,0.40000000000000002]
[null,null,18,0.42199999999999999]
[null,null,19,0.41399999999999998]
[null,null,20,0.39300000000000002]
[]
[null,null,"PROMEDIO",{"$f":"AVERAGE(D14:D33)","v":0.40154999999999996}]
[null,null,"DS",{"$f":"STDEV(D14:D33)","v":0.022772270764890211}]
[]
[null,null,"Criterio de Evaluación","Precisión Intraensayo CV%","Aceptado: Si/No"]
[]
[null,null,"% CV: < 9.9",{"$f":"D36/D35*100","v":5.6710922089130156},"SI"]
[]
["COMENTARIOS: "]
[]
["Realizado por:---------------------------------------",null,null,null,"FECHA:…………………………."]
[]
["REVISADO………………………………………………",null,null,null,"FECHA:…………………………."]

{"sheet":"INTRAENSAYO CP-HBsAg II","index":1,"n_rows":46,"n_cols":5,"rows_kept":46,"cut":false,"hidden":false,"merged":[[37,3,38,3],[37,2,38,2],[37,4,38,4],[2,1,2,4],[3,1,3,4]]}
["NOTICE: This document is an example only. It must be revised to reflect your lab’s specific processes and/or specific protocol requirements."]
["Approved and current. Effective starting 10-Aug-2023. Last reviewed on 10-Aug-2023.\r\nRDP 508 (version 1.0). Precision (AngloLab HbsAg)_Example. Page 2 of 4"]
[null,"ANGLOLAB S.A"]
[null,"CLINICA ANGLO AMERICANA"]
["           PRUEBA DE PRECISION INTRAENSAYO COBAS PURE E402"]
[]
["PRUEBA: HBsAg II",null,"Fecha: 11 de enero 2023"]
[]
["CONTROL : PC HBsAg II 2",null,"Lote: 55190299",null,"Expira: 28-02-2023"]
["MARCA REACTIVO: ROCHE",null,"Lote kit: 64979002",null,"Expira: 31-05-2024"]
["MARCA CALIBRADOR: ROCHE",null,"Lote kit: 64979002",null,"Expira: 31-05-2024"]
[]
[null,null,"CORRIDA","PC HBsAg II 2      (COI)"]
[null,null,1,4.1500000000000004]
[null,null,2,4.0499999999999998]
[null,null,3,4.1200000000000001]
[null,null,4,4.04]
[null,null,5,4.0599999999999996]
[null,null,6,4.0199999999999996]
[null,null,7,4.0800000000000001]
[null,null,8,4.1399999999999997]
[null,null,9,4.1699999999999999]
[null,null,10,4.0700000000000003]
[null,null,11,4]
[null,null,12,4.04]
[null,null,13,4.0800000000000001]
[null,null,14,3.9900000000000002]
[null,null,15,4.0599999999999996]
[null,null,16,4.0999999999999996]
[null,null,17,4.0700000000000003]
[null,null,18,4.1699999999999999]
[null,null,19,4.1500000000000004]
[null,null,20,4.0999999999999996]
[]
[null,null,"PROMEDIO",{"$f":"AVERAGE(D14:D34)","v":4.0830000000000002}]
[null,null,"DS",{"$f":"STDEV(D14:D34)","v":0.053812540115517817}]
[]
[null,null,"Criterio de Evaluación","Precisión Intraensayo CV%","Aceptado: Si/No"]
[]
[null,null,"%CV: < 2.2",{"$f":"D36/D35*100","v":1.3179657143158905},"SI"]
[]
["COMENTARIOS: "]
[]
["Realizado por:---------------------------------------",null,null,null,"FECHA:…………………………."]
[]
["REVISADO………………………………………………",null,null,null,"FECHA:…………………………."]

{"sheet":"INTERENSAYO CN HBsAg","index":2,"n_rows":46,"n_cols":5,"rows_kept":46,"cut":false,"hidden":false,"merged":[[37,3,38,3],[37,2,38,2],[37,4,38,4],[3,1,3,4],[2,1,2,4]]}
["NOTICE: This document is an example only. It must be revised to reflect your lab’s specific processes and/or specific protocol requirements."]
["Approved and current. Effective starting 10-Aug-2023. Last reviewed on 10-Aug-2023.\r\nRDP 508 (version 1.0). Precision (AngloLab HbsAg)_Example. Page 3 of 4"]
[null,"ANGLOLAB S.A"]
[null,"CLINICA ANGLO AMERICANA"]
["           PRUEBA DE PRECISION INTERENSAYO COBAS PURE E402"]
[]
["PRUEBA: HBsAg II",null,"Fecha: Del 03 al 22 de enero 2023"]
[]
["CONTROL : PC HBsAg II 1",null,"Lote: 55190199",null,"Expira: 28-02-2023"]
["MARCA REACTIVO: ROCHE",null,"Lote kit: 64979002",null,"Expira: 31-05-2024"]
["MARCA CALIBRADOR: ROCHE",null,"Lote kit: 64979002",null,"Expira: 31-05-2024"]
[]
[null,null,"CORRIDA","PC HBsAg II 1      (COI)"]
[null,null,1,0.40799999999999997]
[null,null,2,0.40300000000000002]
[null,null,3,0.40300000000000002]
[null,null,4,0.39700000000000002]
[null,null,5,0.378]
[null,null,6,0.45200000000000001]
[null,null,7,0.435]
[null,null,8,0.46400000000000002]
[null,null,9,0.41899999999999998]
[null,null,10,0.45100000000000001]
[null,null,11,0.40400000000000003]
[null,null,12,0.34899999999999998]
[null,null,13,0.40500000000000003]
[null,null,14,0.35699999999999998]
[null,null,15,0.38800000000000001]
[null,null,16,0.38100000000000001]
[null,null,17,0.40899999999999997]
[null,null,18,0.38700000000000001]
[null,null,19,0.42299999999999999]
[null,null,20,0.41299999999999998]
[]
[null,null,"PROMEDIO",{"$f":"AVERAGE(D14:D33)","v":0.40630000000000005}]
[null,null,"DS",{"$f":"STDEV(D14:D33)","v":0.029657873721563567}]
[]
[null,null,"Criterio de Evaluación","Precisión Intraensayo CV%","Aceptado: Si/No"]
[]
[null,null,"%CV: ≤ 11.0",{"$f":"(D36/D35)*100","v":7.2995012851497814},"SI"]
[]
["COMENTARIOS:"]
[]
["Realizado por:---------------------------------------",null,null,null,"FECHA:…………………………."]
[]
["REVISADO………………………………………………",null,null,null,"FECHA:…………………………."]

{"sheet":"INTERENSAYO CP HBsAg","index":3,"n_rows":46,"n_cols":5,"rows_kept":46,"cut":false,"hidden":false,"merged":[[37,3,38,3],[37,2,38,2],[37,4,38,4],[2,1,2,4],[3,1,3,4]]}
["NOTICE: This document is an example only. It must be revised to reflect your lab’s specific processes and/or specific protocol requirements."]
["Approved and current. Effective starting 10-Aug-2023. Last reviewed on 10-Aug-2023.\r\nRDP 508 (version 1.0). Precision (AngloLab HbsAg)_Example. Page 4 of 4"]
[null,"ANGLOLAB S.A"]
[null,"CLINICA ANGLO AMERICANA"]
["           PRUEBA DE PRECISION INTERENSAYO COBAS PURE E402"]
[]
["PRUEBA: HBsAg II",null,"Fecha: Del 03 al 22 de enero 2023"]
[]
["CONTROL :PC HBsAg II 2",null,"Lote: 55190299",null,"Expira: 28-02-2023"]
["MARCA REACTIVO: ROCHE",null,"Lote kit: 64979002",null,"Expira: 31-05-2024"]
["MARCA CALIBRADOR: ROCHE",null,"Lote kit: 64979002",null,"Expira: 31-05-2024"]
[]
[null,null,"CORRIDA","PC HBsAg II 2      (COI)"]
[null,null,1,4.1699999999999999]
[null,null,2,4.0999999999999996]
[null,null,3,4.1200000000000001]
[null,null,4,4.0499999999999998]
[null,null,5,4.0199999999999996]
[null,null,6,3.96]
[null,null,7,3.9900000000000002]
[null,null,8,4.1600000000000001]
[null,null,9,3.9500000000000002]
[null,null,10,4.1299999999999999]
[null,null,11,4.0800000000000001]
[null,null,12,4.1100000000000003]
[null,null,13,4.04]
[null,null,14,4.0599999999999996]
[null,null,15,4.0499999999999998]
[null,null,16,4.0300000000000002]
[null,null,17,3.96]
[null,null,18,4.0099999999999998]
[null,null,19,3.9399999999999999]
[null,null,20,4.1500000000000004]
[]
[null,null,"PROMEDIO",{"$f":"AVERAGE(D14:D33)","v":4.0540000000000003}]
[null,null,"DS",{"$f":"STDEV(D14:D33)","v":0.072358755335371922}]
[]
[null,null,"Criterio de Evaluación","Precisión Intraensayo CV%","Aceptado: Si/No"]
[]
[null,null,"%CV ≤ 3.0",{"$f":"(D36/D35)*100","v":1.7848730965804618},"SI"]
[]
["COMENTARIOS: "]
[]
["Realizado por:---------------------------------------",null,null,null,"FECHA:…………………………."]
[]
["REVISADO………………………………………………",null,null,null,"FECHA:…………………………."]