{"text_format":"2","workbook":"xlsx","n_sheets":4,"sheets":["INTRAENSAYO PC U1 HCG-BETA","INTRAENSAYO PC U2 HCG-BETA","INTERENSAYO PC U1 HCG-BETA","INTERENSAYO PC U2 HCG-BETA"],"not_rendered":{"charts":0,"images":0,"macros":0,"external_links":0,"defined_names":0,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"run_cost":"one","rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"INTRAENSAYO PC U1 HCG-BETA","index":0,"n_rows":46,"n_cols":5,"rows_kept":46,"cut":false,"hidden":false,"merged":[[37,3,38,3],[37,2,38,2],[37,4,38,4],[3,1,3,4],[2,1,2,4]]}
["NOTICE: This document is an example only. It must be revised to reflect your lab’s specific processes and/or specific protocol requirements."]
["Approved and current. Effective starting 10-Aug-2023. Last reviewed on 10-Aug-2023.\r\nRDP 509 (version 1.0). Precision (AngloLab HCG)_Example. Page 1 of 4"]
[null,"ANGLOLAB S.A"]
[null,"CLINICA ANGLO AMERICANA"]
["           PRUEBA DE PRECISION INTRAENSAYO COBAS PURE E402"]
[]
["PRUEBA: HCG-BETA",null,"Fecha: 11 de enero 2023"]
[]
["CONTROL : PC U1",null,"Lote: 55801290",null,"Expira: 30-09-2023"]
["MARCA REACTIVO: ROCHE",null,"Lote kit: 58457502",null,"Expira: 31-07-2023"]
["MARCA CALIBRADOR: ROCHE",null,"Lote kit: 59546502",null,"Expira: 31-05-2023"]
[]
[null,null,"CORRIDA","PC U1"]
[null,null,1,5.0700000000000003]
[null,null,2,5.04]
[null,null,3,5.2199999999999998]
[null,null,4,5.1399999999999997]
[null,null,5,5.1100000000000003]
[null,null,6,5.0599999999999996]
[null,null,7,5.0199999999999996]
[null,null,8,5.04]
[null,null,9,5.1699999999999999]
[null,null,10,5.0599999999999996]
[null,null,11,5.2000000000000002]
[null,null,12,5]
[null,null,13,5.0700000000000003]
[null,null,14,5.0899999999999999]
[null,null,15,5.1399999999999997]
[null,null,16,5.1299999999999999]
[null,null,17,5.1500000000000004]
[null,null,18,5.1100000000000003]
[null,null,19,5.0700000000000003]
[null,null,20,5.0700000000000003]
[]
[null,null,"PROMEDIO",{"$f":"AVERAGE(D14:D33)","v":5.0980000000000008}]
[null,null,"DS",{"$f":"STDEV(D14:D33)","v":0.059169693522773528}]
[]
[null,null,"Criterio de Evaluación","Precisión Intraensayo CV%","Aceptado: Si/No"]
[]
[null,null,"%CV: < 2.5",{"$f":"D36/D35*100","v":1.1606452240638194},"SI"]
[]
["COMENTARIOS: "]
[]
["Realizado por:---------------------------------------",null,null,null,"FECHA:…………………………."]
[]
["REVISADO………………………………………………",null,null,null,"FECHA:…………………………."]

{"sheet":"INTRAENSAYO PC U2 HCG-BETA","index":1,"n_rows":46,"n_cols":5,"rows_kept":46,"cut":false,"hidden":false,"merged":[[37,3,38,3],[37,2,38,2],[37,4,38,4],[3,1,3,4],[2,1,2,4]]}
["NOTICE: This document is an example only. It must be revised to reflect your lab’s specific processes and/or specific protocol requirements."]
["Approved and current. Effective starting 10-Aug-2023. Last reviewed on 10-Aug-2023.\r\nRDP 509 (version 1.0). Precision (AngloLab HCG)_Example. Page 2 of 4"]
[null,"ANGLOLAB S.A"]
[null,"CLINICA ANGLO AMERICANA"]
["           PRUEBA DE PRECISION INTRAENSAYO COBAS PURE E402"]
[]
["PRUEBA: HCG-BETA",null,"Fecha: 11 de enero 2023"]
[]
["CONTROL : PC U2",null,"Lote: 55801490",null,"Expira: 30-09-2023"]
["MARCA REACTIVO: ROCHE",null,"Lote kit: 58457502",null,"Expira: 31-07-2023"]
["MARCA CALIBRADOR: ROCHE",null,"Lote kit: 59546502",null,"Expira: 31-05-2023"]
[]
[null,null,"CORRIDA","PC U2"]
[null,null,1,39.5]
[null,null,2,39.700000000000003]
[null,null,3,39.5]
[null,null,4,39.799999999999997]
[null,null,5,39.600000000000001]
[null,null,6,39.200000000000003]
[null,null,7,39.799999999999997]
[null,null,8,38.899999999999999]
[null,null,9,39.299999999999997]
[null,null,10,39.299999999999997]
[null,null,11,39.799999999999997]
[null,null,12,39.100000000000001]
[null,null,13,39.700000000000003]
[null,null,14,38.700000000000003]
[null,null,15,39.600000000000001]
[null,null,16,39.200000000000003]
[null,null,17,39.100000000000001]
[null,null,18,39.5]
[null,null,19,39.600000000000001]
[null,null,20,39.700000000000003]
[]
[null,null,"PROMEDIO",{"$f":"AVERAGE(D14:D33)","v":39.430000000000014}]
[null,null,"DS",{"$f":"STDEV(D14:D33)","v":0.31639415790733938}]
[]
[null,null,"Criterio de Evaluación","Precisión Intraensayo CV%","Aceptado: Si/No"]
[]
[null,null,"%CV: < 2.1",{"$f":"D36/D35*100","v":0.80241987803027959},"SI"]
[]
["COMENTARIOS: "]
[]
["Realizado por:---------------------------------------",null,null,null,"FECHA:…………………………."]
[]
["REVISADO………………………………………………",null,null,null,"FECHA:…………………………."]

{"sheet":"INTERENSAYO PC U1 HCG-BETA","index":2,"n_rows":46,"n_cols":5,"rows_kept":46,"cut":false,"hidden":false,"merged":[[37,3,38,3],[37,2,38,2],[37,4,38,4],[2,1,2,3]]}
["NOTICE: This document is an example only. It must be revised to reflect your lab’s specific processes and/or specific protocol requirements."]
["Approved and current. Effective starting 10-Aug-2023. Last reviewed on 10-Aug-2023.\r\nRDP 509 (version 1.0). Precision (AngloLab HCG)_Example. Page 3 of 4"]
[null,"ANGLOLAB S.A"]
[null,"          CLINICA ANGLO AMERICANA"]
["           PRUEBA DE PRECISION INTERENSAYO COBAS PURE E402"]
[]
["PRUEBA: HCG-BETA",null,"Fecha: del 03 al 24 de enero 2023"]
[]
["CONTROL : PC U1",null,"Lote: 55801290",null,"Expira: 30-09-2023"]
["MARCA REACTIVO: ROCHE",null,"Lote kit: 58457502",null,"Expira: 31-07-2023"]
["MARCA CALIBRADOR: ROCHE",null,"Lote kit: 59546502",null,"Expira: 31-05-2023"]
[]
[null,null,"CORRIDA","PC U1"]
[null,null,1,5.1100000000000003]
[null,null,2,5.29]
[null,null,3,5.21]
[null,null,4,5.1500000000000004]
[null,null,5,5.1799999999999997]
[null,null,6,5.0499999999999998]
[null,null,7,5.1900000000000004]
[null,null,8,5.0499999999999998]
[null,null,9,5.2300000000000004]
[null,null,10,5.0899999999999999]
[null,null,11,5.0700000000000003]
[null,null,12,4.8200000000000003]
[null,null,13,5.25]
[null,null,14,5.1100000000000003]
[null,null,15,5.0800000000000001]
[null,null,16,5.1799999999999997]
[null,null,17,5.0599999999999996]
[null,null,18,5.1900000000000004]
[null,null,19,5.4199999999999999]
[null,null,20,5.0499999999999998]
[]
[null,null,"PROMEDIO",{"$f":"AVERAGE(D14:D33)","v":5.1390000000000002}]
[null,null,"DS",{"$f":"STDEV(D14:D33)","v":0.12143440941295711}]
[]
[null,null,"Criterio de Evaluación","Precisión Intraensayo CV%","Aceptado: Si/No"]
[]
[null,null,"%CV: < 4.0",{"$f":"(D36/D35)*100","v":2.3629968751305137},"SI"]
[]
["COMENTARIOS: "]
[]
["Realizado por:---------------------------------------",null,null,null,"FECHA:…………………………."]
[]
["REVISADO………………………………………………",null,null,null,"FECHA:…………………………."]

{"sheet":"INTERENSAYO PC U2 HCG-BETA","index":3,"n_rows":46,"n_cols":5,"rows_kept":46,"cut":false,"hidden":false,"merged":[[37,3,38,3],[37,2,38,2],[37,4,38,4],[2,1,2,3]]}
["NOTICE: This document is an example only. It must be revised to reflect your lab’s specific processes and/or specific protocol requirements."]
["Approved and current. Effective starting 10-Aug-2023. Last reviewed on 10-Aug-2023.\r\nRDP 509 (version 1.0). Precision (AngloLab HCG)_Example. Page 4 of 4"]
[null,"ANGLOLAB S.A"]
[null,"          CLINICA ANGLO AMERICANA"]
["           PRUEBA DE PRECISION INTERENSAYO COBAS E402"]
[]
["PRUEBA: HCG-BETA",null,"Fecha: del 03 al 24 de enero 2023"]
[]
["CONTROL : PC U2",null,"Lote: 55801490",null,"Expira: 30-09-2023"]
["MARCA REACTIVO: ROCHE",null,"Lote kit: 58457502",null,"Expira: 31-07-2023"]
["MARCA CALIBRADOR: ROCHE",null,"Lote kit: 59546502",null,"Expira: 31-05-2023"]
[]
[null,null,"CORRIDA","PC U2"]
[null,null,1,40.399999999999999]
[null,null,2,41]
[null,null,3,39.799999999999997]
[null,null,4,39.700000000000003]
[null,null,5,41.200000000000003]
[null,null,6,39.100000000000001]
[null,null,7,39.799999999999997]
[null,null,8,39]
[null,null,9,39.899999999999999]
[null,null,10,39.399999999999999]
[null,null,11,39.100000000000001]
[null,null,12,39.299999999999997]
[null,null,13,40.399999999999999]
[null,null,14,39.799999999999997]
[null,null,15,39.5]
[null,null,16,40.399999999999999]
[null,null,17,39.5]
[null,null,18,40.700000000000003]
[null,null,19,39.899999999999999]
[null,null,20,38.700000000000003]
[]
[null,null,"PROMEDIO",{"$f":"AVERAGE(D14:D33)","v":39.829999999999998}]
[null,null,"DS",{"$f":"STDEV(D14:D33)","v":0.67753422908289096}]
[]
[null,null,"Criterio de Evaluación","Precisión Intraensayo CV%","Aceptado: Si/No"]
[]
[null,null,"%CV: < 3.5",{"$f":"(D36/D35)*100","v":1.7010650993795908},"SI"]
[]
["COMENTARIOS: "]
[]
["Realizado por:---------------------------------------",null,null,null,"FECHA:…………………………."]
[]
["REVISADO………………………………………………",null,null,null,"FECHA:…………………………."]