{"text_format":"2","workbook":"xlsx","n_sheets":4,"sheets":["20 RUNS PV-1","20 RUNS PV-2","4 RUNS PV-1 ","4 RUNS PV-2"],"not_rendered":{"charts":0,"images":0,"macros":0,"external_links":0,"defined_names":0,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"run_cost":"one","rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"20 RUNS PV-1","index":0,"n_rows":31,"n_cols":26,"rows_kept":31,"cut":false,"hidden":false,"merged":[]}
["NOTICE: This document is an example only. It must be revised to reflect your lab’s specific processes and/or specific protocol requirements."]
["Approved and current. Effective starting 10-Aug-2023. Last reviewed on 10-Aug-2023.\r\nRDP 510 (version 1.0). Precision (Chemistry)_Example. Page 1 of 4"]
[null,null,null,"Vitros 250",null,"Serial Number: 123456"]
[null,null,null,"Date:  3 August 2023"]
[null,null,null,"Within Run Precision--Precinormal"]
[]
[null,"URIC","ALT","ALB","AMYL","AST","BU","BC","TBIL","CA","Cl","CHOL","DHDL","CK","CREA","ECO2","ALKP","PHOS","GLU","LAC","LIPA","K","TP","NA","TRIG","BUN"]
[null,3.7999999999999998,40,2.3399999999999999,82,29,0.72999999999999998,0.51000000000000001,1.2,9.0399999999999991,79.900000000000006,146,35.399999999999999,189,0.90000000000000002,22.600000000000001,102,3.7799999999999998,75,1.3100000000000001,134,3.2000000000000002,3.9900000000000002,119.8,120.40000000000001,21]
[null,3.8100000000000001,43,2.3300000000000001,83,29,0.71999999999999997,0.57999999999999996,1.22,8.9299999999999997,79.900000000000006,144,36.100000000000001,192,0.90000000000000002,22,101,3.6899999999999999,75,1.23,132,3.1099999999999999,3.9900000000000002,120.09999999999999,119.7,20.899999999999999]
[null,3.8100000000000001,43,2.2999999999999998,84,29,0.71999999999999997,0.60999999999999999,1.2,8.8699999999999992,79.5,145,35.799999999999997,182,0.90000000000000002,22.199999999999999,101,3.6699999999999999,74,1.22,131,3.1000000000000001,4.0499999999999998,119.90000000000001,120.59999999999999,21.100000000000001]
[null,3.8199999999999998,46,2.3199999999999998,83,29,0.70999999999999996,0.56000000000000005,1.23,9.0500000000000007,79.099999999999994,146,36,190,0.90000000000000002,22,101,3.6899999999999999,74,1.22,135,3.1000000000000001,3.9700000000000002,120.3,120.7,20.699999999999999]
[null,3.7799999999999998,44,2.3300000000000001,84,29,0.70999999999999996,0.54000000000000004,1.21,9.0899999999999999,79.599999999999994,145,37,193,0.90000000000000002,21.800000000000001,103,3.7200000000000002,75,1.23,132,3.0600000000000001,3.98,120.40000000000001,121.7,21]
[null,3.8100000000000001,44,2.3100000000000001,84,29,0.69999999999999996,0.60999999999999999,1.1799999999999999,8.8300000000000001,79.900000000000006,145,35.700000000000003,190,0.90000000000000002,22.800000000000001,103,3.6899999999999999,75,1.23,132,3.1000000000000001,3.98,120.3,121.09999999999999,20.800000000000001]
[null,3.77,42,2.3300000000000001,87,30,0.71999999999999997,0.57999999999999996,1.21,8.8100000000000005,79.599999999999994,144,36.200000000000003,184,0.90000000000000002,21.600000000000001,103,3.6600000000000001,74,1.23,133,3.0600000000000001,3.9399999999999999,120,121.90000000000001,20.699999999999999]
[null,3.7599999999999998,44,2.3300000000000001,85,29,0.69999999999999996,0.58999999999999997,1.2,8.8699999999999992,79.799999999999997,145,37.100000000000001,187,0.90000000000000002,22.399999999999999,101,3.6600000000000001,74,1.22,129,3.1000000000000001,4,120.09999999999999,121.90000000000001,20.899999999999999]
[null,3.79,41,2.3300000000000001,84,29,0.70999999999999996,0.56999999999999995,1.22,8.9700000000000006,79.599999999999994,145,36.5,188,0.90000000000000002,22,102,3.7000000000000002,75,1.22,133,3.1000000000000001,3.9700000000000002,120,121,20.899999999999999]
[null,3.7799999999999998,43,2.3199999999999998,82,29,0.70999999999999996,0.57999999999999996,1.21,9.1400000000000006,79.700000000000003,146,36.100000000000001,184,0.90000000000000002,22.399999999999999,102,3.71,75,1.23,131,3.1299999999999999,4.0199999999999996,120.09999999999999,121.5,20.600000000000001]
[null,3.7799999999999998,44,2.2999999999999998,86,30,0.68999999999999995,0.56999999999999995,1.2,9.1199999999999992,79.599999999999994,147,36.600000000000001,193,0.90000000000000002,23.100000000000001,103,3.6699999999999999,75,1.29,133,3.0600000000000001,3.96,120.5,121.40000000000001,20.600000000000001]
[null,3.7799999999999998,45,2.3199999999999998,86,29,0.71999999999999997,0.55000000000000004,1.1899999999999999,9.0399999999999991,79.799999999999997,146,35.899999999999999,190,0.90000000000000002,21.899999999999999,102,3.6699999999999999,73.599999999999994,1.23,131,3.0899999999999999,3.9500000000000002,120.40000000000001,121,20.800000000000001]
[null,3.8399999999999999,45,2.3300000000000001,85,29,0.70999999999999996,0.54000000000000004,1.1899999999999999,8.8599999999999994,79.799999999999997,142,36.399999999999999,184,0.90000000000000002,22.300000000000001,103,3.6899999999999999,72.900000000000006,1.23,131,3.0699999999999998,3.9399999999999999,120.2,120.3,20.800000000000001]
[null,3.8100000000000001,43,2.3399999999999999,86,29,0.70999999999999996,0.57999999999999996,1.21,8.8900000000000006,79.5,145,36.100000000000001,187,0.90000000000000002,22.600000000000001,103,3.6899999999999999,72.799999999999997,1.21,133,3.1000000000000001,3.9500000000000002,120.59999999999999,122.09999999999999,20.600000000000001]
[null,3.8100000000000001,43,2.2999999999999998,84,29,0.70999999999999996,0.60999999999999999,1.21,8.9199999999999999,79.599999999999994,147,36.600000000000001,187,0.90000000000000002,22.300000000000001,101,3.6699999999999999,74,1.23,134,3.1099999999999999,3.9700000000000002,120.40000000000001,120.3,20.899999999999999]
[null,3.7999999999999998,44,2.3100000000000001,83,29,0.69999999999999996,0.58999999999999997,1.21,9.1600000000000001,79.5,144,37.100000000000001,188,0.90000000000000002,22.199999999999999,103,3.6899999999999999,73.799999999999997,1.22,133,3.0699999999999998,4.0099999999999998,120.40000000000001,121.59999999999999,20.800000000000001]
[null,3.77,46,2.3399999999999999,84,29,0.68000000000000005,0.59999999999999998,1.21,9.5199999999999996,79.799999999999997,143,36,186,0.90000000000000002,22.600000000000001,103,3.6899999999999999,73.599999999999994,1.21,132,3.1000000000000001,4.0199999999999996,120.59999999999999,121.90000000000001,20.600000000000001]
[null,3.8199999999999998,42,2.3199999999999998,86,30,0.68000000000000005,0.58999999999999997,1.1899999999999999,8.9199999999999999,79.700000000000003,144,36.299999999999997,185,0.90000000000000002,22.600000000000001,103,3.6699999999999999,73.700000000000003,1.24,131,3.0899999999999999,3.98,119.8,121.40000000000001,20.899999999999999]
[null,3.8100000000000001,44,2.3100000000000001,85,29,0.68999999999999995,0.62,1.1899999999999999,8.7599999999999998,79.299999999999997,145,35.899999999999999,177,0.90000000000000002,22.5,102,3.6600000000000001,73.299999999999997,1.22,130,3.1200000000000001,3.9900000000000002,120.3,121.5,20.800000000000001]
[null,3.7799999999999998,42,2.3500000000000001,86,29,0.68000000000000005,0.58999999999999997,1.23,8.8200000000000003,80.099999999999994,143,36.299999999999997,186,0.90000000000000002,23,103,3.7000000000000002,73.5,1.22,130,3.0800000000000001,4,120.40000000000001,120.90000000000001,20.399999999999999]
["Mean",{"$f":"AVERAGE(B8:B27)","v":3.7964999999999995},{"$f":"AVERAGE(C8:C27)","v":43.399999999999999},{"$f":"AVERAGE(D8:D27)","v":2.3230000000000004},{"$f":"AVERAGE(E8:E27)","v":84.450000000000003},{"$f":"AVERAGE(F8:F27)","v":29.149999999999999},{"$f":"AVERAGE(G8:G27)","v":0.70499999999999996},{"$f":"AVERAGE(H8:H27)","v":0.5784999999999999},{"$f":"AVERAGE(I8:I27)","v":1.2055000000000002},{"$f":"AVERAGE(J8:J27)","v":8.9804999999999993},{"$f":"AVERAGE(K8:K27)","v":79.664999999999992},{"$f":"AVERAGE(L8:L27)","v":144.84999999999999},{"$f":"AVERAGE(M8:M27)","v":36.254999999999995},{"$f":"AVERAGE(N8:N27)","v":187.09999999999999},{"$f":"AVERAGE(O8:O27)","v":0.90000000000000002},{"$f":"AVERAGE(P8:P27)","v":22.345000000000006},{"$f":"AVERAGE(Q8:Q27)","v":102.25},{"$f":"AVERAGE(R8:R27)","v":3.6884999999999999},{"$f":"AVERAGE(S8:S27)","v":74.109999999999985},{"$f":"AVERAGE(T8:T27)","v":1.232},{"$f":"AVERAGE(U8:U27)","v":132},{"$f":"AVERAGE(V8:V27)","v":3.0975000000000006},{"$f":"AVERAGE(W8:W27)","v":3.9829999999999997},{"$f":"AVERAGE(X8:X27)","v":120.23000000000002},{"$f":"AVERAGE(Y8:Y27)","v":121.14500000000001},{"$f":"AVERAGE(Z8:Z27)","v":20.789999999999999}]
["SD",{"$f":"STDEV(B8:B27)","v":0.020844032340469238},{"$f":"STDEV(C8:C27)","v":1.5355437918998287},{"$f":"STDEV(D8:D27)","v":0.014545753585442803},{"$f":"STDEV(E8:E27)","v":1.4317821063276353},{"$f":"STDEV(F8:F27)","v":0.36634754853252322},{"$f":"STDEV(G8:G27)","v":0.014689774459950363},{"$f":"STDEV(H8:H27)","v":0.027772572611727626},{"$f":"STDEV(I8:I27)","v":0.013562719801760006},{"$f":"STDEV(J8:J27)","v":0.174007713077686},{"$f":"STDEV(K8:K27)","v":0.22774639635487595},{"$f":"STDEV(L8:L27)","v":1.3088765773505315},{"$f":"STDEV(M8:M27)","v":0.45937600698063685},{"$f":"STDEV(N8:N27)","v":3.8919620271853406},{"$f":"STDEV(O8:O27)","v":0},{"$f":"STDEV(P8:P27)","v":0.39400106866005247},{"$f":"STDEV(Q8:Q27)","v":0.85069630922340067},{"$f":"STDEV(R8:R27)","v":0.027390932037715244},{"$f":"STDEV(S8:S27)","v":0.74685304691365551},{"$f":"STDEV(T8:T27)","v":0.024623480450043717},{"$f":"STDEV(U8:U27)","v":1.5217718205053643},{"$f":"STDEV(V8:V27)","v":0.031267100584278208},{"$f":"STDEV(W8:W27)","v":0.02867237805201878},{"$f":"STDEV(X8:X27)","v":0.24516374764205051},{"$f":"STDEV(Y8:Y27)","v":0.65652594449075241},{"$f":"STDEV(Z8:Z27)","v":0.17137217117324363}]
["CV",{"$f":"B29/B28","v":0.0054903285501038427},{"$f":"C29/C28","v":0.035381193361747207},{"$f":"D29/D28","v":0.0062616244448742151},{"$f":"E29/E28","v":0.016954199009208231},{"$f":"F29/F28","v":0.012567668903345566},{"$f":"G29/G28","v":0.020836559517660093},{"$f":"H29/H28","v":0.048007904255363235},{"$f":"I29/I28","v":0.011250700789514727},{"$f":"J29/J28","v":0.019376172048069264},{"$f":"K29/K28","v":0.0028588011843956062},{"$f":"L29/L28","v":0.0090360826879567254},{"$f":"M29/M28","v":0.012670693889963782},{"$f":"N29/N28","v":0.020801507360691293},{"$f":"O29/O28","v":0},{"$f":"P29/P28","v":0.01763262782099138},{"$f":"Q29/Q28","v":0.0083197683053633318},{"$f":"R29/R28","v":0.0074260355260174172},{"$f":"S29/S28","v":0.010077628483519844},{"$f":"T29/T28","v":0.019986591274386133},{"$f":"U29/U28","v":0.011528574397767912},{"$f":"V29/V28","v":0.010094302044964714},{"$f":"W29/W28","v":0.0071986889410039622},{"$f":"X29/X28","v":0.0020391229114368335},{"$f":"Y29/Y28","v":0.0054193400015745789},{"$f":"Z29/Z28","v":0.0082430096764426961}]
["n",{"$f":"COUNT(B8:B27)","v":20},{"$f":"COUNT(C8:C27)","v":20},{"$f":"COUNT(D8:D27)","v":20},{"$f":"COUNT(E8:E27)","v":20},{"$f":"COUNT(F8:F27)","v":20},{"$f":"COUNT(G8:G27)","v":20},{"$f":"COUNT(H8:H27)","v":20},{"$f":"COUNT(I8:I27)","v":20},{"$f":"COUNT(J8:J27)","v":20},{"$f":"COUNT(K8:K27)","v":20},{"$f":"COUNT(L8:L27)","v":20},{"$f":"COUNT(M8:M27)","v":20},{"$f":"COUNT(N8:N27)","v":20},{"$f":"COUNT(O8:O27)","v":20},{"$f":"COUNT(P8:P27)","v":20},{"$f":"COUNT(Q8:Q27)","v":20},{"$f":"COUNT(R8:R27)","v":20},{"$f":"COUNT(S8:S27)","v":20},{"$f":"COUNT(T8:T27)","v":20},{"$f":"COUNT(U8:U27)","v":20},{"$f":"COUNT(V8:V27)","v":20},{"$f":"COUNT(W8:W27)","v":20},{"$f":"COUNT(X8:X27)","v":20},{"$f":"COUNT(Y8:Y27)","v":20},{"$f":"COUNT(Z8:Z27)","v":20}]

{"sheet":"20 RUNS PV-2","index":1,"n_rows":31,"n_cols":26,"rows_kept":31,"cut":false,"hidden":false,"merged":[]}
["NOTICE: This document is an example only. It must be revised to reflect your lab’s specific processes and/or specific protocol requirements."]
["Approved and current. Effective starting 10-Aug-2023. Last reviewed on 10-Aug-2023.\r\nRDP 510 (version 1.0). Precision (Chemistry)_Example. Page 2 of 4"]
[{"$nulls":4},"Vitros 250",null,"Serial Number: 123456"]
[{"$nulls":4},"Date:  3 August 2023"]
[{"$nulls":4},"Within Run Precision--Preciabnormal"]
[]
[null,"URIC","ALT","ALB","AMYL","AST","BU","BC","TBIL","CA","Cl","CHOL","DHDL","CK","CREA","ECO2","ALKP","PHOS","GLU","LAC","LIPA","K","TP","NA","TRIG","BUN"]
[null,10.300000000000001,183,4.3600000000000003,319,150,8.4000000000000004,5.4900000000000002,12.76,11.44,112.90000000000001,233,50.700000000000003,801,5.5499999999999998,15.1,488,6.7300000000000004,275.10000000000002,3.48,618,5.2599999999999998,7.1600000000000001,145.69999999999999,250,54.5]
[null,10.380000000000001,186,4.3499999999999996,314,149,8.3000000000000007,5.4400000000000004,12.630000000000001,11.460000000000001,112.90000000000001,229,50.200000000000003,806,5.5,15.199999999999999,490,6.6699999999999999,275.10000000000002,3.5,618,5.2400000000000002,7.0899999999999999,145.80000000000001,249,54]
[null,10.359999999999999,183,4.3799999999999999,316,150,8.4100000000000001,5.4000000000000004,12.49,11.369999999999999,112,229,49.299999999999997,790,5.4199999999999999,14.699999999999999,494,6.7300000000000004,274.5,3.52,617,5.2699999999999996,7.1100000000000003,145.69999999999999,249,53.799999999999997]
[null,10.33,188,4.4199999999999999,315,150,8.4600000000000009,5.4299999999999997,12.75,11.609999999999999,112.3,227,51.5,796,5.5300000000000002,14.800000000000001,482,6.7000000000000002,272.89999999999998,3.5,617,5.2300000000000004,7.0599999999999996,145.30000000000001,248,53.600000000000001]
[null,10.279999999999999,184,4.46,322,152,8.4900000000000002,5.4500000000000002,12.859999999999999,11.43,112.5,226,50.100000000000001,791,5.5700000000000003,15.1,495,6.7999999999999998,273.10000000000002,3.4900000000000002,611,5.2699999999999996,7.04,145,250,53.799999999999997]
[null,10.24,186,4.4500000000000002,318,150,8.4399999999999995,5.4900000000000002,12.73,11.369999999999999,111.90000000000001,226,50.299999999999997,796,5.5499999999999998,14.9,489,6.71,273.19999999999999,3.4900000000000002,609,5.21,7.1200000000000001,145.59999999999999,250,54.299999999999997]
[null,10.35,183,4.3600000000000003,321,150,8.3599999999999994,5.4699999999999998,12.720000000000001,11.31,112.59999999999999,226,49.5,791,5.5099999999999998,15.1,487,6.7400000000000002,277.39999999999998,3.4900000000000002,624,5.2599999999999998,7.0899999999999999,145.59999999999999,253,54]
[null,10.5,185,4.4800000000000004,317,150,8.3900000000000006,5.4699999999999998,12.73,11.32,112.8,231,49.399999999999999,785,5.5099999999999998,15.1,473,6.75,274.10000000000002,3.46,619,5.2400000000000002,7.1500000000000004,145.40000000000001,251,53.5]
[null,10.41,185,4.4299999999999997,322,149,8.5199999999999996,5.3700000000000001,12.75,11.390000000000001,112.7,223,49.899999999999999,790,5.54,15.300000000000001,489,6.7199999999999998,274.5,3.5099999999999998,616,5.2599999999999998,7.1500000000000004,145.80000000000001,248,55.299999999999997]
[null,10.4,187,4.4000000000000004,320,149,8.4299999999999997,5.4500000000000002,12.84,11.800000000000001,112.3,227,49.200000000000003,778,5.54,15,491,6.75,273,3.52,615,5.25,7.1600000000000001,145.59999999999999,249,55.200000000000003]
[null,10.380000000000001,184,4.4400000000000004,314,151,8.4299999999999997,5.4000000000000004,12.699999999999999,11.51,112.7,230,49.399999999999999,807,5.5,14.5,480,6.6600000000000001,274.30000000000001,3.4900000000000002,607,5.2599999999999998,7.0899999999999999,145.69999999999999,250,54.899999999999999]
[null,10.359999999999999,187,4.4299999999999997,319,151,8.3499999999999996,5.3700000000000001,12.76,11.43,112.09999999999999,232,49.100000000000001,798,5.4900000000000002,15.199999999999999,485,6.7599999999999998,274.19999999999999,3.48,616,5.2300000000000004,7.0499999999999998,145.5,250,54.5]
[null,10.390000000000001,183,4.4100000000000001,327,150,8.4700000000000006,5.5300000000000002,12.83,11.390000000000001,112.40000000000001,229,50.5,799,5.5499999999999998,15.199999999999999,486,6.7300000000000004,275.60000000000002,3.4900000000000002,616,5.25,7.0999999999999996,145.59999999999999,251,54.100000000000001]
[null,10.289999999999999,185,4.4500000000000002,317,150,8.4299999999999997,5.46,12.77,11.380000000000001,112.09999999999999,230,49.899999999999999,790,5.4000000000000004,14.9,486,6.75,276.69999999999999,3.4700000000000002,619,5.2599999999999998,7.1200000000000001,145.59999999999999,252,54]
[null,10.390000000000001,185,4.4400000000000004,312,150,8.4199999999999999,5.3899999999999997,12.77,11.369999999999999,112.09999999999999,229,50.399999999999999,796,5.54,15.300000000000001,490,6.7300000000000004,273.30000000000001,3.5,608,5.25,7.1200000000000001,146,248,54.100000000000001]
[null,10.35,186,4.46,314,150,8.1799999999999997,5.5999999999999996,12.869999999999999,11.68,112.7,231,50.399999999999999,783,5.5999999999999996,14,489,6.75,272,3.6499999999999999,618,5.2599999999999998,7.1399999999999997,145.5,249,53.799999999999997]
[null,10.380000000000001,186,4.4400000000000004,313,149,8.1199999999999992,5.5599999999999996,12.84,11.5,112,228,48.5,778,5.5300000000000002,14.9,481,6.6799999999999997,274,3.4700000000000002,621,5.25,7.0599999999999996,145.80000000000001,249,55.299999999999997]
[null,10.380000000000001,184,4.4000000000000004,309,151,8.1300000000000008,5.6399999999999997,12.800000000000001,11.32,112.7,230,49.299999999999997,796,5.5700000000000003,15.699999999999999,484,6.7199999999999998,273.10000000000002,3.4900000000000002,610,5.2800000000000002,7.1399999999999997,145,251,54.899999999999999]
[null,10.460000000000001,184,4.4100000000000001,315,150,8.0899999999999999,5.5599999999999996,12.76,11.41,112.59999999999999,231,49.799999999999997,779,5.5800000000000001,15.300000000000001,476,6.7599999999999998,280,3.4900000000000002,619,5.2599999999999998,7.04,145.5,251,54.700000000000003]
[null,10.48,183,4.5,314,151,8.1699999999999999,5.6100000000000003,12.630000000000001,11.4,112.8,232,50.100000000000001,787,5.5499999999999998,14.300000000000001,489,6.7999999999999998,278.19999999999999,3.4700000000000002,620,5.25,7.2000000000000002,145.09999999999999,252,54.700000000000003]
["Mean",{"$f":"AVERAGE(B8:B27)","v":10.3705},{"$f":"AVERAGE(C8:C27)","v":184.84999999999999},{"$f":"AVERAGE(D8:D27)","v":4.4234999999999989},{"$f":"AVERAGE(E8:E27)","v":316.89999999999998},{"$f":"AVERAGE(F8:F27)","v":150.09999999999999},{"$f":"AVERAGE(G8:G27)","v":8.349499999999999},{"$f":"AVERAGE(H8:H27)","v":5.4789999999999992},{"$f":"AVERAGE(I8:I27)","v":12.749500000000001},{"$f":"AVERAGE(J8:J27)","v":11.4445},{"$f":"AVERAGE(K8:K27)","v":112.45500000000001},{"$f":"AVERAGE(L8:L27)","v":228.94999999999999},{"$f":"AVERAGE(M8:M27)","v":49.874999999999986},{"$f":"AVERAGE(N8:N27)","v":791.85000000000002},{"$f":"AVERAGE(O8:O27)","v":5.5265000000000004},{"$f":"AVERAGE(P8:P27)","v":14.98},{"$f":"AVERAGE(Q8:Q27)","v":486.19999999999999},{"$f":"AVERAGE(R8:R27)","v":6.7320000000000011},{"$f":"AVERAGE(S8:S27)","v":274.71499999999997},{"$f":"AVERAGE(T8:T27)","v":3.4980000000000002},{"$f":"AVERAGE(U8:U27)","v":615.89999999999998},{"$f":"AVERAGE(V8:V27)","v":5.2520000000000007},{"$f":"AVERAGE(W8:W27)","v":7.1094999999999997},{"$f":"AVERAGE(X8:X27)","v":145.53999999999999},{"$f":"AVERAGE(Y8:Y27)","v":250},{"$f":"AVERAGE(Z8:Z27)","v":54.350000000000001}]
["SD",{"$f":"STDEV(B8:B27)","v":0.064764756497732265},{"$f":"STDEV(C8:C27)","v":1.5312533566021211},{"$f":"STDEV(D8:D27)","v":0.04029823033857758},{"$f":"STDEV(E8:E27)","v":4.191470034800628},{"$f":"STDEV(F8:F27)","v":0.78806892565241216},{"$f":"STDEV(G8:G27)","v":0.13527651992162304},{"$f":"STDEV(H8:H27)","v":0.080779100997069295},{"$f":"STDEV(I8:I27)","v":0.089822925218102312},{"$f":"STDEV(J8:J27)","v":0.12487783503973204},{"$f":"STDEV(K8:K27)","v":0.33003189638515207},{"$f":"STDEV(L8:L27)","v":2.5230516193063131},{"$f":"STDEV(M8:M27)","v":0.68277683801247691},{"$f":"STDEV(N8:N27)","v":8.5241144006746818},{"$f":"STDEV(O8:O27)","v":0.04891184336183442},{"$f":"STDEV(P8:P27)","v":0.3847076812334268},{"$f":"STDEV(Q8:Q27)","v":5.5875802125409315},{"$f":"STDEV(R8:R27)","v":0.036792447737631066},{"$f":"STDEV(S8:S27)","v":2.0110942295178482},{"$f":"STDEV(T8:T27)","v":0.039148838794367376},{"$f":"STDEV(U8:U27)","v":4.5986267973225035},{"$f":"STDEV(V8:V27)","v":0.016091841672756044},{"$f":"STDEV(W8:W27)","v":0.04477722635447634},{"$f":"STDEV(X8:X27)","v":0.26832815729997489},{"$f":"STDEV(Y8:Y27)","v":1.4142135623730951},{"$f":"STDEV(Z8:Z27)","v":0.56707931311979731}]
["CV",{"$f":"B29/B28","v":0.0062450948843095575},{"$f":"C29/C28","v":0.0082837617343907016},{"$f":"D29/D28","v":0.0091100328560139235},{"$f":"E29/E28","v":0.013226475338594598},{"$f":"F29/F28","v":0.0052502926425876898},{"$f":"G29/G28","v":0.016201750993667052},{"$f":"H29/H28","v":0.014743402262651818},{"$f":"I29/I28","v":0.0070452115940313189},{"$f":"J29/J28","v":0.010911602519964354},{"$f":"K29/K28","v":0.0029347907730661333},{"$f":"L29/L28","v":0.011020098795834519},{"$f":"M29/M28","v":0.013689761163157435},{"$f":"N29/N28","v":0.010764809497600154},{"$f":"O29/O28","v":0.0088504194991105428},{"$f":"P29/P28","v":0.0256814206430859},{"$f":"Q29/Q28","v":0.011492349264790069},{"$f":"R29/R28","v":0.0054653071505690824},{"$f":"S29/S28","v":0.0073206567880088396},{"$f":"T29/T28","v":0.011191777814284555},{"$f":"U29/U28","v":0.0074665153390526115},{"$f":"V29/V28","v":0.0030639454822460094},{"$f":"W29/W28","v":0.0062982243975633085},{"$f":"X29/X28","v":0.0018436729235947157},{"$f":"Y29/Y28","v":0.005656854249492381},{"$f":"Z29/Z28","v":0.010433842007723961}]
["n",{"$f":"COUNT(B8:B27)","v":20},{"$f":"COUNT(C8:C27)","v":20},{"$f":"COUNT(D8:D27)","v":20},{"$f":"COUNT(E8:E27)","v":20},{"$f":"COUNT(F8:F27)","v":20},{"$f":"COUNT(G8:G27)","v":20},{"$f":"COUNT(H8:H27)","v":20},{"$f":"COUNT(I8:I27)","v":20},{"$f":"COUNT(J8:J27)","v":20},{"$f":"COUNT(K8:K27)","v":20},{"$f":"COUNT(L8:L27)","v":20},{"$f":"COUNT(M8:M27)","v":20},{"$f":"COUNT(N8:N27)","v":20},{"$f":"COUNT(O8:O27)","v":20},{"$f":"COUNT(P8:P27)","v":20},{"$f":"COUNT(Q8:Q27)","v":20},{"$f":"COUNT(R8:R27)","v":20},{"$f":"COUNT(S8:S27)","v":20},{"$f":"COUNT(T8:T27)","v":20},{"$f":"COUNT(U8:U27)","v":20},{"$f":"COUNT(V8:V27)","v":20},{"$f":"COUNT(W8:W27)","v":20},{"$f":"COUNT(X8:X27)","v":20},{"$f":"COUNT(Y8:Y27)","v":20},{"$f":"COUNT(Z8:Z27)","v":20}]

{"sheet":"4 RUNS PV-1 ","index":2,"n_rows":31,"n_cols":26,"rows_kept":31,"cut":false,"hidden":false,"merged":[]}
["NOTICE: This document is an example only. It must be revised to reflect your lab’s specific processes and/or specific protocol requirements."]
["Approved and current. Effective starting 10-Aug-2023. Last reviewed on 10-Aug-2023.\r\nRDP 510 (version 1.0). Precision (Chemistry)_Example. Page 3 of 4"]
[{"$nulls":4},"Vitros 250",null,"Serial Number: 123456"]
[{"$nulls":4},"Date:  3-7 August 2023"]
[{"$nulls":4},"Between Day Precision--Precinormal"]
[]
[null,"URIC","ALT","ALB","AMYL","AST","BU","BC","TBIL","CA","Cl","CHOL","DHDL","CK","CREA","ECO2","ALKP","PHOS","GLU","LAC","LIPA","K","TP","NA","TRIG","BUN"]
["4 RUNS-A",3.8199999999999998,45,2.3799999999999999,83,29,0.68999999999999995,0.58999999999999997,1.22,8.9399999999999995,81.700000000000003,154,35,197,0.89000000000000001,21.5,107,3.71,74.799999999999997,1.3,130,3.0699999999999998,3.9900000000000002,117.59999999999999,122.09999999999999,20.800000000000001]
["2023-08-03",3.7999999999999998,45,2.3900000000000001,81,29,0.68999999999999995,0.57999999999999996,1.1499999999999999,8.9499999999999993,81.700000000000003,150,35.399999999999999,187,0.88,21.800000000000001,107,3.73,74,1.27,133,2.9399999999999999,3.9199999999999999,117.40000000000001,120.5,20.600000000000001]
[null,3.7999999999999998,45,2.3799999999999999,87,30,0.69999999999999996,0.53000000000000003,1.1599999999999999,8.9399999999999995,81.299999999999997,151,35.200000000000003,184,0.89000000000000001,22.100000000000001,106,3.7000000000000002,74,1.25,132,3.0099999999999998,4,117.5,121.8,20.5]
[null,3.8100000000000001,46,2.3799999999999999,87,29,0.70999999999999996,0.51000000000000001,1.1599999999999999,8.8699999999999992,81.599999999999994,150,36.200000000000003,187,0.89000000000000001,22.899999999999999,105,3.6699999999999999,73.5,1.25,135,3,3.96,117.3,121.90000000000001,20.600000000000001]
["4 RUNS-B",3.8999999999999999,44,2.4199999999999999,82,29,0.68999999999999995,0.56000000000000005,1.1799999999999999,8.9800000000000004,81,150,34.700000000000003,187,0.89000000000000001,21.699999999999999,101,3.71,75.5,1.21,125,3.1099999999999999,3.9500000000000002,119.3,119.2,20.899999999999999]
["2023-08-04",3.8700000000000001,41,2.4100000000000001,86,28,0.69999999999999996,0.56000000000000005,1.1599999999999999,8.9800000000000004,81.200000000000003,149,35.399999999999999,185,0.89000000000000001,22.800000000000001,101,3.6499999999999999,75.400000000000006,1.21,126,3.0699999999999998,3.9700000000000002,119.7,118.2,20.800000000000001]
[null,3.8399999999999999,44,2.3799999999999999,86,29,0.71999999999999997,0.55000000000000004,1.1599999999999999,9.0099999999999998,81.5,150,35.100000000000001,186,0.90000000000000002,21.699999999999999,99,3.6299999999999999,74.299999999999997,1.1899999999999999,123,3.1299999999999999,3.9900000000000002,119.2,117.5,20.800000000000001]
[null,3.8199999999999998,43,2.3700000000000001,85,29,0.69999999999999996,0.56000000000000005,1.1699999999999999,8.9299999999999997,81.200000000000003,150,35.100000000000001,190,0.88,21.300000000000001,101,3.6400000000000001,74.299999999999997,1.1899999999999999,127,3.0699999999999998,4.0300000000000002,118.90000000000001,118.90000000000001,20.800000000000001]
["4 RUNS-C",3.8900000000000001,45,2.3300000000000001,85,29,0.68999999999999995,0.56999999999999995,1.1899999999999999,9.2699999999999996,81.5,152,34.899999999999999,189,0.89000000000000001,21.800000000000001,103,3.6499999999999999,75.5,1.3,130,3.1200000000000001,4,119.7,120.5,21.199999999999999]
["2023-08-05",3.8900000000000001,43,2.3399999999999999,87,29,0.68999999999999995,0.57999999999999996,1.1899999999999999,9.2100000000000009,82,151,34.799999999999997,184,0.90000000000000002,23.100000000000001,102,3.6499999999999999,75.900000000000006,1.3,133,3.0800000000000001,3.9399999999999999,120.40000000000001,118.3,21.300000000000001]
[null,3.8900000000000001,43,2.3300000000000001,86,29,0.68000000000000005,0.57999999999999996,1.1699999999999999,9.1400000000000006,81.799999999999997,150,35.200000000000003,187,0.90000000000000002,23.899999999999999,103,3.6299999999999999,75.5,1.29,132,3.0800000000000001,3.9399999999999999,119.8,118.59999999999999,21.100000000000001]
[null,3.8999999999999999,42,2.3500000000000001,86,30,0.68000000000000005,0.56000000000000005,1.1799999999999999,9.1799999999999997,81.900000000000006,153,34.5,189,0.93999999999999995,22.5,100,3.6699999999999999,74.099999999999994,1.29,130,3.0600000000000001,3.98,119.8,119.59999999999999,21.100000000000001]
["4 RUNS-D",3.8999999999999999,45,2.3999999999999999,81,30,0.71999999999999997,0.57999999999999996,1.1699999999999999,8.8000000000000007,81.299999999999997,148,33.799999999999997,189,0.90000000000000002,22.199999999999999,103,3.6800000000000002,74.900000000000006,1.3300000000000001,136,3.02,4.0300000000000002,119.90000000000001,120.5,20.800000000000001]
["2023-08-06",3.9300000000000002,42,2.3700000000000001,83,30,0.72999999999999998,0.56999999999999995,1.1499999999999999,8.7899999999999991,81.299999999999997,148,33.799999999999997,186,0.89000000000000001,22.800000000000001,102,3.7400000000000002,75,1.3200000000000001,133,3.04,3.98,119.7,119.3,20.899999999999999]
[null,3.9100000000000001,43,2.3599999999999999,83,30,0.70999999999999996,0.62,1.1599999999999999,8.8300000000000001,80.700000000000003,148,35.399999999999999,188,0.89000000000000001,22.600000000000001,100,3.6899999999999999,74.099999999999994,1.3100000000000001,136,3.04,3.98,119.3,118.5,20.800000000000001]
[null,3.9100000000000001,44,2.3700000000000001,83,30,0.70999999999999996,0.60999999999999999,1.1699999999999999,8.7400000000000002,81.299999999999997,147,35.100000000000001,192,0.90000000000000002,22.899999999999999,101,3.71,74.400000000000006,1.3,137,3.0600000000000001,4.0099999999999998,119.8,119.5,20.5]
["4 RUNS-E",3.9399999999999999,44,2.3700000000000001,null,29,0.68999999999999995,0.60999999999999999,1.1599999999999999,8.9900000000000002,81.200000000000003,149,35.200000000000003,195,0.91000000000000003,23.199999999999999,103,3.6499999999999999,75.900000000000006,1.3300000000000001,136,3.0099999999999998,4.0599999999999996,120.40000000000001,123.2,21.300000000000001]
["2023-08-07",3.9100000000000001,41,2.3799999999999999,null,29,0.69999999999999996,0.59999999999999998,1.1699999999999999,8.9600000000000009,81.400000000000006,147,35.399999999999999,194,0.90000000000000002,21.899999999999999,103,3.6699999999999999,75.700000000000003,1.3300000000000001,139,3.0099999999999998,4,120.3,121.8,21.100000000000001]
[null,3.8900000000000001,43,2.3599999999999999,null,29,0.68999999999999995,0.58999999999999997,1.1399999999999999,9.0199999999999996,81.5,147,35.600000000000001,194,0.89000000000000001,22.899999999999999,104,3.6600000000000001,75.299999999999997,1.3200000000000001,139,2.98,4,120.40000000000001,121.09999999999999,21]
[null,3.9100000000000001,41,2.3799999999999999,null,29,0.68000000000000005,0.59999999999999998,1.1499999999999999,9.0299999999999994,82.099999999999994,145,35.299999999999997,193,0.93999999999999995,22.600000000000001,105,3.6400000000000001,76,1.3200000000000001,137,3.0499999999999998,4.04,120.5,121.09999999999999,20.899999999999999]
["Mean",{"$f":"AVERAGE(B8:B27)","v":3.8764999999999992},{"$f":"AVERAGE(C8:C27)","v":43.450000000000003},{"$f":"AVERAGE(D8:D27)","v":2.3724999999999996},{"$f":"AVERAGE(E8:E27)","v":84.4375},{"$f":"AVERAGE(F8:F27)","v":29.25},{"$f":"AVERAGE(G8:G27)","v":0.6984999999999999},{"$f":"AVERAGE(H8:H27)","v":0.5754999999999999},{"$f":"AVERAGE(I8:I27)","v":1.1679999999999997},{"$f":"AVERAGE(J8:J27)","v":8.9780000000000015},{"$f":"AVERAGE(K8:K27)","v":81.460000000000008},{"$f":"AVERAGE(L8:L27)","v":149.44999999999999},{"$f":"AVERAGE(M8:M27)","v":35.055},{"$f":"AVERAGE(N8:N27)","v":189.15000000000001},{"$f":"AVERAGE(O8:O27)","v":0.89800000000000024},{"$f":"AVERAGE(P8:P27)","v":22.41},{"$f":"AVERAGE(Q8:Q27)","v":102.8},{"$f":"AVERAGE(R8:R27)","v":3.6739999999999995},{"$f":"AVERAGE(S8:S27)","v":74.905000000000001},{"$f":"AVERAGE(T8:T27)","v":1.2805},{"$f":"AVERAGE(U8:U27)","v":132.44999999999999},{"$f":"AVERAGE(V8:V27)","v":3.0474999999999999},{"$f":"AVERAGE(W8:W27)","v":3.9884999999999997},{"$f":"AVERAGE(X8:X27)","v":119.34500000000003},{"$f":"AVERAGE(Y8:Y27)","v":120.10499999999998},{"$f":"AVERAGE(Z8:Z27)","v":20.890000000000001}]
["SD",{"$f":"STDEV(B8:B27)","v":0.04451788522900841},{"$f":"STDEV(C8:C27)","v":1.5035046776746235},{"$f":"STDEV(D8:D27)","v":0.023591925469711467},{"$f":"STDEV(E8:E27)","v":2.1281838892977896},{"$f":"STDEV(F8:F27)","v":0.5501196042201808},{"$f":"STDEV(G8:G27)","v":0.014244112357114603},{"$f":"STDEV(H8:H27)","v":0.027236778540480094},{"$f":"STDEV(I8:I27)","v":0.017947291248483575},{"$f":"STDEV(J8:J27)","v":0.14054780044713019},{"$f":"STDEV(K8:K27)","v":0.34397062298063152},{"$f":"STDEV(L8:L27)","v":2.1878853044122666},{"$f":"STDEV(M8:M27)","v":0.55675280350486545},{"$f":"STDEV(N8:N27)","v":3.8151360602073869},{"$f":"STDEV(O8:O27)","v":0.016091841672756169},{"$f":"STDEV(P8:P27)","v":0.67426373100395076},{"$f":"STDEV(Q8:Q27)","v":2.3078812338319508},{"$f":"STDEV(R8:R27)","v":0.033308762874212666},{"$f":"STDEV(S8:S27)","v":0.76741191848418766},{"$f":"STDEV(T8:T27)","v":0.04751454071068531},{"$f":"STDEV(U8:U27)","v":4.6280040030915677},{"$f":"STDEV(V8:V27)","v":0.047779647726224471},{"$f":"STDEV(W8:W27)","v":0.035876249290552852},{"$f":"STDEV(X8:X27)","v":1.0644865330042323},{"$f":"STDEV(Y8:Y27)","v":1.573288543816824},{"$f":"STDEV(Z8:Z27)","v":0.24039441275580956}]
["CV",{"$f":"B29/B28","v":0.011484041075456834},{"$f":"C29/C28","v":0.034603099601257153},{"$f":"D29/D28","v":0.0099439095762745931},{"$f":"E29/E28","v":0.025204250354377965},{"$f":"F29/F28","v":0.018807507836587377},{"$f":"G29/G28","v":0.020392430003027352},{"$f":"H29/H28","v":0.047327156456090529},{"$f":"I29/I28","v":0.015365831548359229},{"$f":"J29/J28","v":0.015654689290168208},{"$f":"K29/K28","v":0.0042225708689004604},{"$f":"L29/L28","v":0.014639580491216237},{"$f":"M29/M28","v":0.015882265112105703},{"$f":"N29/N28","v":0.020169897225521476},{"$f":"O29/O28","v":0.017919645515318668},{"$f":"P29/P28","v":0.030087627443282052},{"$f":"Q29/Q28","v":0.022450206554785513},{"$f":"R29/R28","v":0.0090660759047938681},{"$f":"S29/S28","v":0.010245136085497465},{"$f":"T29/T28","v":0.037106240305103719},{"$f":"U29/U28","v":0.034941517577135282},{"$f":"V29/V28","v":0.01567830934412616},{"$f":"W29/W28","v":0.0089949227254739507},{"$f":"X29/X28","v":0.0089194062005465841},{"$f":"Y29/Y28","v":0.013099275998641392},{"$f":"Z29/Z28","v":0.011507631055807063}]
["n",{"$f":"COUNT(B8:B27)","v":20},{"$f":"COUNT(C8:C27)","v":20},{"$f":"COUNT(D8:D27)","v":20},{"$f":"COUNT(E8:E27)","v":16},{"$f":"COUNT(F8:F27)","v":20},{"$f":"COUNT(G8:G27)","v":20},{"$f":"COUNT(H8:H27)","v":20},{"$f":"COUNT(I8:I27)","v":20},{"$f":"COUNT(J8:J27)","v":20},{"$f":"COUNT(K8:K27)","v":20},{"$f":"COUNT(L8:L27)","v":20},{"$f":"COUNT(M8:M27)","v":20},{"$f":"COUNT(N8:N27)","v":20},{"$f":"COUNT(O8:O27)","v":20},{"$f":"COUNT(P8:P27)","v":20},{"$f":"COUNT(Q8:Q27)","v":20},{"$f":"COUNT(R8:R27)","v":20},{"$f":"COUNT(S8:S27)","v":20},{"$f":"COUNT(T8:T27)","v":20},{"$f":"COUNT(U8:U27)","v":20},{"$f":"COUNT(V8:V27)","v":20},{"$f":"COUNT(W8:W27)","v":20},{"$f":"COUNT(X8:X27)","v":20},{"$f":"COUNT(Y8:Y27)","v":20},{"$f":"COUNT(Z8:Z27)","v":20}]

{"sheet":"4 RUNS PV-2","index":3,"n_rows":31,"n_cols":26,"rows_kept":31,"cut":false,"hidden":false,"merged":[]}
["NOTICE: This document is an example only. It must be revised to reflect your lab’s specific processes and/or specific protocol requirements."]
["Approved and current. Effective starting 10-Aug-2023. Last reviewed on 10-Aug-2023.\r\nRDP 510 (version 1.0). Precision (Chemistry)_Example. Page 4 of 4"]
[{"$nulls":5},"Vitros 250",null,"Serial Number: 123456"]
[{"$nulls":5},"Date:  3-7 August 2023"]
[{"$nulls":5},"Between Day Precision--Preciabnormal"]
[]
[null,"URIC","ALT","ALB","AMYL","AST","BU","BC","TBIL","CA","Cl","CHOL","DHDL","CK","CREA","ECO2","ALKP","PHOS","GLU","LAC","LIPA","K","TP","NA","TRIG","BUN"]
["4 RUNS-A",10.279999999999999,184,4.5099999999999998,319,154,8.5399999999999991,5.5700000000000003,12.880000000000001,11.49,114.40000000000001,237,49.899999999999999,802,5.46,15,504,6.6200000000000001,273.30000000000001,3.6000000000000001,618,5.46,6.9699999999999998,148,250.30000000000001,54.5]
["2023-08-03",10.31,185,4.54,321,153,8.3699999999999992,5.4299999999999997,12.720000000000001,11.550000000000001,114.90000000000001,234,49.899999999999999,797,5.4699999999999998,15,495,6.71,273,3.4900000000000002,621,5.3499999999999996,2.9500000000000002,148.59999999999999,248.69999999999999,54]
[null,10.369999999999999,183,4.5499999999999998,316,154,8.4499999999999993,5.4400000000000004,12.75,11.550000000000001,114.8,233,50.600000000000001,792,5.46,15,500,6.6699999999999999,272,3.52,628,5.4299999999999997,6.9900000000000002,148,251,53.799999999999997]
[null,10.369999999999999,187,4.5099999999999998,326,154,8.4000000000000004,5.4000000000000004,12.609999999999999,11.880000000000001,114.40000000000001,241,49.600000000000001,781,5.5,14,495,6.6799999999999997,270.10000000000002,3.48,627,5.3600000000000003,6.9100000000000001,147.59999999999999,250.59999999999999,53.600000000000001]
["4 RUNS-B",10.31,186,4.4900000000000002,312,149,8.4900000000000002,5.1699999999999999,12.720000000000001,11.44,112.7,236,48.299999999999997,792,5.5199999999999996,14,467,6.79,282.80000000000001,3.6099999999999999,613,5.4299999999999997,7.0700000000000003,148.40000000000001,246.30000000000001,53.799999999999997]
["2023-08-04",10.31,183,4.5300000000000002,310,150,8.5199999999999996,5.1799999999999997,12.74,11.42,112.59999999999999,234,48.600000000000001,770,5.5300000000000002,14.300000000000001,462,6.79,280,3.5699999999999998,627,5.3799999999999999,7.0800000000000001,148.09999999999999,245.5,54.299999999999997]
[null,10.35,184,4.5499999999999998,309,151,8.5600000000000005,5.1399999999999997,12.75,11.390000000000001,112.8,238,49,795,5.5300000000000002,14.199999999999999,454,6.7400000000000002,276,3.5699999999999998,631,5.4400000000000004,7.0599999999999996,148,244.19999999999999,54]
[null,10.26,184,4.5099999999999998,310,151,8.6199999999999992,5.1500000000000004,12.630000000000001,11.35,112,234,49,783,5.4299999999999997,14.800000000000001,465,6.7699999999999996,276.19999999999999,3.5600000000000001,621,5.4299999999999997,7.0300000000000002,148.19999999999999,244.90000000000001,53.5]
["4 RUNS-C",10.42,183,4.5199999999999996,321,150,8.7100000000000009,5.0300000000000002,13.029999999999999,11.32,112.7,236,49.200000000000003,797,5.4500000000000002,14.800000000000001,475,6.7199999999999998,272.89999999999998,3.5299999999999998,618,5.4299999999999997,7.1299999999999999,148.5,253.40000000000001,55.299999999999997]
["2023-08-05",10.449999999999999,182,4.5499999999999998,317,151,8.6699999999999999,5.04,12.890000000000001,11.34,112.5,234,48.5,772,5.4199999999999999,14.6,479,6.79,274.39999999999998,3.4900000000000002,622,5.4100000000000001,7.0800000000000001,148.09999999999999,250.5,55.200000000000003]
[null,10.4,183,4.5099999999999998,313,152,8.7799999999999994,5.1100000000000003,13,11.48,112.2,231,48.700000000000003,791,5.4100000000000001,15,467,6.71,272,3.4900000000000002,618,5.4500000000000002,7.0199999999999996,148.30000000000001,251.80000000000001,54.899999999999999]
[null,10.470000000000001,182,4.54,312,152,8.7899999999999991,5.1100000000000003,12.949999999999999,11.470000000000001,112.90000000000001,233,48.600000000000001,805,5.4900000000000002,15,471,6.75,271.80000000000001,3.4700000000000002,621,5.3399999999999999,7.1299999999999999,148,251.90000000000001,54.5]
["4 RUNS-D",10.32,183,4.5199999999999996,306,152,8.5399999999999991,5.2300000000000004,12.68,11.449999999999999,114,228,48,817,5.46,14.1,467,6.7999999999999998,273.60000000000002,3.7000000000000002,658,5.2000000000000002,7.0499999999999998,147.5,249.40000000000001,54.100000000000001]
["2023-08-06",10.529999999999999,181,4.5300000000000002,309,152,8.6099999999999994,5.2400000000000002,12.65,11.42,112.8,234,48.200000000000003,824,5.5099999999999998,14.5,467,6.7999999999999998,273.60000000000002,3.71,649,5.3099999999999996,7.04,148.5,250.80000000000001,54]
[null,10.470000000000001,182,4.5099999999999998,310,152,8.5099999999999998,5.1600000000000001,12.73,11.44,112.90000000000001,235,50.200000000000003,821,5.46,12.800000000000001,457,6.7699999999999996,273.89999999999998,3.7000000000000002,656,5.2199999999999998,7.1200000000000001,149,250.90000000000001,54.100000000000001]
[null,10.44,181,4.5300000000000002,311,152,8.4800000000000004,5.1900000000000004,12.5,11.380000000000001,113.5,235,49.5,825,5.5300000000000002,13.9,465,6.6799999999999997,273.89999999999998,3.71,652,5.2800000000000002,7.1100000000000003,148.5,251.90000000000001,53.799999999999997]
["4 RUNS-E",10.539999999999999,178,4.5599999999999996,null,152,7.8700000000000001,5.4800000000000004,12.68,11.68,113.8,231,49.600000000000001,826,5.5300000000000002,14.9,477,6.7199999999999998,277.19999999999999,3.7000000000000002,653,5.21,7.2000000000000002,148.59999999999999,254.5,55.299999999999997]
["2023-08-07",10.49,174,4.5599999999999996,null,153,7.8399999999999999,5.5199999999999996,12.720000000000001,11.56,114.8,234,49.799999999999997,822,5.4500000000000002,14.6,479,6.7400000000000002,276,3.7400000000000002,655,5.29,7.1900000000000004,148.5,252.09999999999999,54.899999999999999]
[null,10.460000000000001,176,4.5,null,152,7.8799999999999999,5.5099999999999998,12.74,11.550000000000001,113.5,234,49.600000000000001,813,5.4299999999999997,15,484,6.7400000000000002,273.69999999999999,3.7400000000000002,661,5.21,7.1299999999999999,148.80000000000001,252.09999999999999,54.700000000000003]
[null,10.5,177,4.5599999999999996,null,153,7.9400000000000004,5.4199999999999999,12.710000000000001,11.57,113.59999999999999,231,50,834,5.5800000000000001,14.6,484,6.7300000000000004,277,3.7400000000000002,656,5.3200000000000003,7.0999999999999996,148.69999999999999,254.5,54.700000000000003]
["Mean",{"$f":"AVERAGE(B8:B27)","v":10.4025},{"$f":"AVERAGE(C8:C27)","v":181.90000000000001},{"$f":"AVERAGE(D8:D27)","v":4.5290000000000008},{"$f":"AVERAGE(E8:E27)","v":313.875},{"$f":"AVERAGE(F8:F27)","v":151.94999999999999},{"$f":"AVERAGE(G8:G27)","v":8.4284999999999997},{"$f":"AVERAGE(H8:H27)","v":5.2759999999999998},{"$f":"AVERAGE(I8:I27)","v":12.754000000000001},{"$f":"AVERAGE(J8:J27)","v":11.486499999999999},{"$f":"AVERAGE(K8:K27)","v":113.39000000000001},{"$f":"AVERAGE(L8:L27)","v":234.15000000000001},{"$f":"AVERAGE(M8:M27)","v":49.240000000000009},{"$f":"AVERAGE(N8:N27)","v":802.95000000000005},{"$f":"AVERAGE(O8:O27)","v":5.4809999999999999},{"$f":"AVERAGE(P8:P27)","v":14.505000000000001},{"$f":"AVERAGE(Q8:Q27)","v":475.69999999999999},{"$f":"AVERAGE(R8:R27)","v":6.7359999999999989},{"$f":"AVERAGE(S8:S27)","v":274.66999999999996},{"$f":"AVERAGE(T8:T27)","v":3.6060000000000003},{"$f":"AVERAGE(U8:U27)","v":635.25},{"$f":"AVERAGE(V8:V27)","v":5.3474999999999993},{"$f":"AVERAGE(W8:W27)","v":6.8680000000000003},{"$f":"AVERAGE(X8:X27)","v":148.29500000000002},{"$f":"AVERAGE(Y8:Y27)","v":250.26500000000004},{"$f":"AVERAGE(Z8:Z27)","v":54.350000000000001}]
["SD",{"$f":"STDEV(B8:B27)","v":0.086260010860677513},{"$f":"STDEV(C8:C27)","v":3.322966334942004},{"$f":"STDEV(D8:D27)","v":0.021496633763403331},{"$f":"STDEV(E8:E27)","v":5.5362442142665635},{"$f":"STDEV(F8:F27)","v":1.3562719801759986},{"$f":"STDEV(G8:G27)","v":0.3015272528487703},{"$f":"STDEV(H8:H27)","v":0.17463947826428319},{"$f":"STDEV(I8:I27)","v":0.13343241932272593},{"$f":"STDEV(J8:J27)","v":0.12978828914813556},{"$f":"STDEV(K8:K27)","v":0.91185178851089976},{"$f":"STDEV(L8:L27)","v":2.7961439613567078},{"$f":"STDEV(M8:M27)","v":0.73513335703910132},{"$f":"STDEV(N8:N27)","v":18.997160452634965},{"$f":"STDEV(O8:O27)","v":0.04587167455321197},{"$f":"STDEV(P8:P27)","v":0.55769733349682926},{"$f":"STDEV(Q8:Q27)","v":14.253716043636885},{"$f":"STDEV(R8:R27)","v":0.048925292135319064},{"$f":"STDEV(S8:S27)","v":2.9725232945835938},{"$f":"STDEV(T8:T27)","v":0.10122408700242097},{"$f":"STDEV(U8:U27)","v":17.161385663597706},{"$f":"STDEV(V8:V27)","v":0.088844926994351078},{"$f":"STDEV(W8:W27)","v":0.92489316595001336},{"$f":"STDEV(X8:X27)","v":0.38590290570502256},{"$f":"STDEV(Y8:Y27)","v":2.9799814552584407},{"$f":"STDEV(Z8:Z27)","v":0.56707931311979731}]
["CV",{"$f":"B29/B28","v":0.0082922384869673176},{"$f":"C29/C28","v":0.018268094199791116},{"$f":"D29/D28","v":0.0047464415463465068},{"$f":"E29/E28","v":0.017638372646010557},{"$f":"F29/F28","v":0.0089257780860546154},{"$f":"G29/G28","v":0.035774723005133811},{"$f":"H29/H28","v":0.033100735076626836},{"$f":"I29/I28","v":0.010462005592184876},{"$f":"J29/J28","v":0.011299202467952428},{"$f":"K29/K28","v":0.0080417302099911776},{"$f":"L29/L28","v":0.011941678246238341},{"$f":"M29/M28","v":0.014929597015416352},{"$f":"N29/N28","v":0.023659207239099526},{"$f":"O29/O28","v":0.0083692163023557694},{"$f":"P29/P28","v":0.038448626921532524},{"$f":"Q29/Q28","v":0.029963666267893388},{"$f":"R29/R28","v":0.0072632559583312162},{"$f":"S29/S28","v":0.010822162211321202},{"$f":"T29/T28","v":0.028071016916922063},{"$f":"U29/U28","v":0.027015168301609927},{"$f":"V29/V28","v":0.016614292098055369},{"$f":"W29/W28","v":0.13466703056930887},{"$f":"X29/X28","v":0.002602265118210476},{"$f":"Y29/Y28","v":0.011907304078710329},{"$f":"Z29/Z28","v":0.010433842007723961}]
["n",{"$f":"COUNT(B8:B27)","v":20},{"$f":"COUNT(C8:C27)","v":20},{"$f":"COUNT(D8:D27)","v":20},{"$f":"COUNT(E8:E27)","v":16},{"$f":"COUNT(F8:F27)","v":20},{"$f":"COUNT(G8:G27)","v":20},{"$f":"COUNT(H8:H27)","v":20},{"$f":"COUNT(I8:I27)","v":20},{"$f":"COUNT(J8:J27)","v":20},{"$f":"COUNT(K8:K27)","v":20},{"$f":"COUNT(L8:L27)","v":20},{"$f":"COUNT(M8:M27)","v":20},{"$f":"COUNT(N8:N27)","v":20},{"$f":"COUNT(O8:O27)","v":20},{"$f":"COUNT(P8:P27)","v":20},{"$f":"COUNT(Q8:Q27)","v":20},{"$f":"COUNT(R8:R27)","v":20},{"$f":"COUNT(S8:S27)","v":20},{"$f":"COUNT(T8:T27)","v":20},{"$f":"COUNT(U8:U27)","v":20},{"$f":"COUNT(V8:V27)","v":20},{"$f":"COUNT(W8:W27)","v":20},{"$f":"COUNT(X8:X27)","v":20},{"$f":"COUNT(Y8:Y27)","v":20},{"$f":"COUNT(Z8:Z27)","v":20}]