Verification Summary Report for GeneXpert (MTB/RIF, MTB/RIF Ultra, MTB XDR)

| (Please fill in the table with your laboratory’s information and details on the method being verified) |  |  |
| --- | --- | --- |
| Date(s) Verification Performed: |  |  |
| Laboratory Name: |  |  |
| City, Country: |  |  |
| Instrument/Method/Reagent being validated: |  |  |
|  | ☐ Primary ☐ Back-up |  |
| (if applicable) Serial Number(s): |  |  |
| Analyte(s): |  |  |
| Kit Name: |  |  |
| Reason for Verification: | ☐ Initial Verification | ☐ Re-verification (choose one below) ☐ Instrument move ☐ Instrument modified ☐ Method change ☐ Other: _____________________ |
| Regulatory Status: (check all that apply) | ☐ FDA Approved ☐ FDA Cleared ☐ CE Marked ☐ EUA ☐ None |  |

Procedure: Refer to the (insert lab name) Verification Plan for (insert test/method name). This summary was written using “VAL 3000_Mycobacteriology Validation Guidelines” as a reference, please refer to this document if more details are needed.

Results:  All raw data reports and statistical analysis can be found in the (insert location details).

- Precision testing for this method is not required for this qualitative method.

- Accuracy- Identify the reference method and materials used to establish accuracy.

  - Accuracy was demonstrated using (insert comparison method details)

  - (Describe sample type used for accuracy testing, for example patient samples or EQA panel details)

  - (Describe number of samples tested for each level, for example 10 positive and 10 negative)

  - The following tables list the accuracy testing results

| (Enter Name of Method Being Validated) | Diagnostic Sensitivity and Specificity (Results from Comparison Study) |  | Total |
| --- | --- | --- | --- |
|  | Positive | Negative |  |
| Positive | # true positive (TP) | # false positive (FP) | TP+FP |
| Negative | # false negative (FN) | # true negative (TN) | FN+TN |
| Total | TP+FN | FP+TN | N |

|  | Lab Result (%) | Expected Result (%) | Acceptability |
| --- | --- | --- | --- |
| Sensitivity= 100 x TP/(TP+FN) | (00.00%) | (Insert manufacturer’s claim) | (Acceptable) |
| Specificity= 100 x TN/(FP+TN) | (00.00%) | (Insert manufacturer’s claim) | (Acceptable) |
| Positive Predictive Value= 100 x TP/(TP+FP) | (00.00%) | (Insert manufacturer’s claim) | (Acceptable) |
| Negative Predictive Value= 100 x TN/(TN+FN) | (00.00%) | (Insert manufacturer’s claim) | (Acceptable) |

- Analytical Sensitivity and Specificity- Refer to test kit package insert. For FDA approved, unmodified tests, manufacturer’s stated sensitivity and specificity will be used.

| Prepared by: |  |
| --- | --- |
| Date: |  |

Method Approval

☐ Approved

☐ Not Approved (provide recommendations/corrective actions below)

Additional comments, if needed:

| Laboratory Director: |  |
| --- | --- |
| Date: |  |