Validation Summary Report for Quantitative Method

| (Please fill in the table with your laboratory’s information and details on the method being validated) |  |  |
| --- | --- | --- |
| Date(s) Validation Performed: |  |  |
| Laboratory Name: |  |  |
| City, Country: |  |  |
| Instrument/Method/Reagent being validated: |  |  |
|  | ☐ Primary ☐ Back-up |  |
| (if applicable) Serial Number(s): |  |  |
| Analyte(s): |  |  |
| Kit Name: |  |  |
| Sample Type(s): |  |  |
| Reason for Validation: | ☐ Initial Validation | ☐ Re-validation (choose one below) ☐ Instrument move ☐ Instrument modified ☐ Method change ☐ Other: _____________________ |
| Regulatory Status: (check all that apply) | ☐ FDA Approved ☐ FDA Cleared ☐ CE Marked ☐ EUA ☐ None |  |

Procedure: Refer to the (insert lab name) Validation Plan for (insert instrument name and serial number). This summary was written using “VAL 2001_Quantitative Validation Guidelines” as a reference, please refer to this document if more details are needed.

Results:  All raw data reports and statistical analysis can be found in the (insert location details).

- Precision- Identify the materials used for short-term and long-term testing

  - Short-Term

(Insert number of samples and materials) were used on the same run.

OR-

(Insert number of samples and materials) were used within the same day.

- (E.g. 20 high- and low-quality control samples were used on the same run.)

| Analyte | Expected Results |  | Observed Results: Short-term |  | Acceptability |
| --- | --- | --- | --- | --- | --- |
|  | Manufacturer’s Precision | 25% of CLIA | Normal Control | Abn Control |  |
|  |  |  | CV% | CV% |  |
| (ALT) | (2.60%) | (5%) | (2.60%) | (4.30%) | (Acceptable) |
| (WBC) | (2.60%) | (3.75%) | (2.60%) | (3.30%) | (Acceptable) |
| (PT) | (3.50%) | (3.75%) | (2.60%) | (4.30%) | (Acceptable) |

  - Long-Term

(insert sample material) were run two times per day for 10 days.

OR-

(insert sample material) were run once per day for 20 days.

OR-

(insert sample material) were run 5 times per day for 4 days.

- (E.g. High- and low-quality controls were run two times per day for 10 days.)

| Analyte | Expected Results |  | Observed Results: Long-term |  | Acceptability |
| --- | --- | --- | --- | --- | --- |
|  | Manufacturer’s Precision | 33% of CLIA | Normal Control | Abn. Control |  |
|  |  |  | CV% | CV% |  |
| (ALT) | (3.30%) | (6.60%) | (3.80%) | (4.30%) | (Acceptable) |
| (WBC) | (3.30%) | (4.95%) | (3.80%) | (4.30%) | (Acceptable) |
| (PT) | (4.50%) | (4.95%) | (3.80%) | (4.30%) | (Acceptable) |

- Accuracy- Identify the reference method and materials used to establish accuracy.

  - (Insert number of samples and materials) were used to correlate with (insert laboratory name) in (location) on the (test method)

- E.g. 20 patient samples were used to correlate with XYZ Laboratory in Baltimore, Maryland on the Cobas Integra (SN # 1234).

OR-

Accuracy was established by using (insert number of samples, proficiency provider, year and panel name) with results compared to the peer means.

- E.g. Accuracy was established by using 15 CAP samples from the 2022 FH9 survey with results compared to the peer means and 5 patient samples or 5 CAP samples from the 2023 FH9 A event

| Analyte | Total Allowable Error | Correlation Coefficient (R) | Linear Regression Statistics |  | Error Index Range | % of Error Indices -1.0 to 1.0 | Acceptability |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Expected >0.975 | Slope | Intercept | Expected -1.0 to1.0 | Expected ≥ 95% |  |
| (ALT) | (0.5U/L or 20%) | (0.997) | (1.002) | (2.19) | (-0.29-1.05) | (95%) | (Acceptable) |
| (WBC) | (0.12 or 15%) | (0.997) | (1.002) | (2.19) | (-0.29-1.05) | (95%) | (Acceptable) |
| (PT) | (15%) | (0.997) | (1.002) | (2.19) | (-0.29-1.05) | (95%) | (Acceptable) |

  - Alternate Comparison- Required only for analytes identified as unacceptable in section 2. a.

Alternate Correlation was performed with (Insert name) Laboratory in (location) on the (test method)

- E.g. Alternate Correlation was performed with XYZ Laboratory in Baltimore, Maryland on the Cobas Integra (SN#1234).

| Analyte | Total Allowable Error | Correlation Coefficient (R) | MDP Error Index Range |  | Acceptability |
| --- | --- | --- | --- | --- | --- |
|  |  | Expected >0.975 | Expected -1.0 to1.0 | 100% of Error Indices: -1.0 to 1.0 |  |
| (ALT) | (0.5U/L or 20%) | (0.997) | (-0.29-1.00) | (yes) | (Acceptable) |
| (WBC) | (0.12 or 15%) | (0.997) | (-0.29-1.00) | (yes) | (Acceptable) |
| (PT) | (15%) | (0.997) | (-0.29-1.00) | (yes) | (Acceptable) |

- Linearity – Identify the method or materials used to establish linearity.

Note for Coagulation: Linearity study is not applicable, skip to next section.

Linearity was established from (insert number of samples) run in duplicate from the (insert provider, year and panel name)

- E.g. Linearity was established from 6 samples run in duplicate from the CAP 2022 LN2-B panel.

| Analyte | Linear Regression Statistics |  | Allowable Systematic Error | Linear Range Verified | Evaluation |
| --- | --- | --- | --- | --- | --- |
|  | Slope | Intercept | 50% of TEa |  |  |
| (ALT) | (0.97) | (0.282) | (10%) | (20-350) | (Linear) |
| (WBC) | (0.97) | (0.282) | (7.50%) | (1.0-25) | (Linear) |

- Analytical Measurement Range (AMR) and Clinical Reportable Range (CRR)-

Note for Coagulation AMR and CRR studies are not applicable, skip to next section.

| Analyte | Mfg’s AMR | Low Value Verified | High Value Verified | Reportable Range | Dilutions | CRR | DAIDS Toxicity Grade 4 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| (ALT) | (5-700 U/L) | (2.5) | (770) | (5-700) | (1:10) | (5-7000) | (>381) |
| (WBC) | (0.5-25 109/uL) | (1.0) | (25) | (1.0-25) | (1:10) | (1.0-250) | (<1.0) |

- Analytical Sensitivity and Specificity- See below for examples of how to enter information from the package insert into the summary form.

Chemistry

| Summary of Manufacturer’s Claims for Sensitivity and Specificity |  |  |
| --- | --- | --- |
| Analyte | Specificity (Interfering Substances) | Sensitivity |
| ALT | Icterus – No significant effect Hemolysis – No significant effect Lipemia – No significant effect; >Abs flagging may occur Other –Calcium dobesilate may cause low results | 5 U/L |
| AST | Icterus – No significant effect Hemolysis – No significant effect Lipemia – No significant effect; >Abs flagging may occur | 5 U/L |
| Albumin | Icterus –No significant effect Hemolysis- No significant effect Lipemia- No significant effect | 0.2 g/dL |

Hematology

| Summary of Manufacturer’s Claims for Sensitivity and Specificity |  |  |
| --- | --- | --- |
| Analyte | Specificity (Interfering Substances) | Sensitivity |
| WBC: | Unlysed Red Cells - False increase Multiple Myeloma - False increase Leukemia - False decrease Chemotherapy - False decrease Cryoglobulins - False increase WBC Agglutination - False decrease PLT Agglutination - False increase | 0 K/uL |
| RBC: | Agglutinated RBC’s - False decrease & falsely elevated MCH, MCHC & MCV Cold Agglutinins - False decrease & falsely elevated MCV | 0 1012/L |
| HGB | High WBC - An extremely high WBC will cause excessive light scatter. In these cases use reference methods. Lipemia - Significant interference - Use a reference method and plasma blank. Turbidity - False increase, abnormal MCH, MCHC values & an increased baseline on the leading edge of the WBC histogram. Fetal Blood - False increase. | 0 g/dL |
| HCT | RBC agglutination - Inaccurate results. | 0 g/dL |
| MCV | RBC agglutination - Inaccurate results. Large Platelets (excessive numbers) - Inaccurate results High WBC’s - Inaccurate results | N/A |
| MCH | Calculated based on HGB & RBC -refer to those for MCH limitations. | N/A |
| MCHC | Calculated based on HBG & HCT - refer to those for MCHC limitations. | N/A |
| RDW | Calculated based on RBC - refer to RBC count for limitations | N/A |
| PLTs | Microcytes, Schistocytes & WBC Fragments - False increase RBC Agglutination - Falsely decrease Giant Platelets (large number) - False decrease Chemotherapy - False decrease Hemolysis - False increase ACD Blood - False decrease Increased Triglycerides and/or Cholesterol - Inaccurate results Platelet Agglutination - False decrease | 0 K/uL |
| MPV | Giant Platelets - Inaccurate results Microcytes, Schistocytes & WBC Fragments - Inaccurate results RBC Agglutination - Inaccurate results Chemotherapy - Inaccurate results | N/A |
| LYM % & Absolute | Erythroblasts - False increase Parasites - False increase Unlysed RBC’s - False increase WBC limitations - Pertains to differential enumeration also | 0 K/uL |
| MON % & Absolute | Large lymphocytes - False increase Atypical Lymphocytes - False increase Blasts - False increase Basophils (excessive numbers) - False increase WBC limitations - Pertains to differential enumeration also | 0 K/uL |
| NEU % & Absolute | Eosinophils (excessive numbers) - Inaccurate results Immature granulocytes - Inaccurate results Plasma cells - Inaccurate results WBC limitations - Pertains to differential enumeration also | 0 K/uL |
| EOS % Absolute | Abnormal granules - Inaccurate results WBC limitations - Pertains to differential enumeration also | 0 K/uL |
| BAS % & Absolute | WBC limitations - Pertains to differential enumeration also | 0 K/uL |

Coagulation

| Summary of Manufacturer’s Claims for Sensitivity and Specificity |  |  |
| --- | --- | --- |
| Analyte | Specificity (Interfering Substances) | Sensitivity |
| Pt | Hemolysis- Can cause decrease times. Lipemia- Can cause increase times Icterus- No significant effect. |  |
| APtt | Hemolysis- Can cause decrease times. Lipemia- Can cause increase times Icterus- No significant effect. |  |
| Fibrinogen | Hemolysis- No significant effect. Lipemia- Can cause increase times Icterus- No significant effect. |  |

- Reference ranges- Describe the methods used to establish or verify reference ranges for each analyte.

- Reference range was transfered/adopted from (insert the source/published range) or established.

- Include information on how “normal” subjects were screened, the total number of subjects included, and any other pertinent information.

- See below for examples of how to enter information for reference ranges into the summary form.

Chemistry

| Analyte | Adult Reference Ranges | Reference Range Cited | % Verified (Expected ≥90%) |
| --- | --- | --- | --- |
| Albumin | 3.4-5.0 mg/dL | Dimension Manufacturer’s range | 100% |
| Alk Phos | 38-126 U/L | Caribbean Reference ranges | 95% |
| ALT | 9-72 U/L | Caribbean Reference ranges | 95% |
| Amylase | 25-125 U/L | Panamanian Reference ranges | 95% |
| AST | 0-37 U/L | Panamanian Reference ranges | 100% |
| BUN | 7-18 mg/dL | Dimension Manufacturer’s range | 100% |
| Cholesterol | 0-200 mg/dL | N/A* | N/A* |
| Creatinine | 0.5-1.4 mg/dL | Panamanian Reference ranges | 95% |
| D. Bilirubin | 0.0-0.3 mg/dL | Dimension Manufacturer’s range | 100% |
| Glucose | 70-110 mg/dL | Dimension Manufacturer’s range | 100% |
| HDL Cholesterol | 40-60 mg/dL | N/A* | N/A* |
| LDL Cholesterol | <130.00 mg/dL | N/A* | N/A* |
| LDH | 100-190 U/L | Dimension Manufacturer’s range | 95% |
| Lipase | 20-300 U/L | Caribbean Reference range | 90% |
| T. Bilirubin | 0.1-1.2 mg/dL | Caribbean Reference range | 95% |
| Total Protein | 6.4-8.2 g/dL | Dimension Manufacturer’s range | 100% |
| Triglycerides | 0-150 mg/dL | N/A* | N/A* |
| Uric Acid | 2.6-7.2 mg/dL | Dimension Manufacturer’s range | 100% |

Hematology

| Analyte | Reference Population | Adult Male | % Verified (Expected >90%) | Adult Female | % Verified (Expected >90%) | Acceptability |
| --- | --- | --- | --- | --- | --- | --- |
| WBC (K/uL) | Ugandan | 3.2 - 8.2 | 90% | 3.6 - 10.3 | 100% | Acceptable |
| Neutrophil % | Ugandan | 21.6 - 66.9 | 100% | 29.3 - 64.0 | 90% | Acceptable |
| Lymphocyte % | Coulter Inst. range | 21.2 - 59.8 | 95% | 26.5 - 59.2 | 90% | Acceptable |
| Monocytes % | Coulter Inst. range | 4.3 - 9.6 | 95% | 3.5 - 10.9 | 95% | Acceptable |
| Eosinophils % | Coulter Inst. range | 1.1 - 28.8 | 95% | 1.0 - 10.9 | 95% | Acceptable |
| Basophils % | Coulter Inst. range | 0.3 - 2.0 | 100% | 0.3 - 1.0 | 100% | Acceptable |
| Neutrophil Abs. | Coulter Inst. range | 1.11 - 3.78 | 95% | 1.21 - 6.01 | 95% | Acceptable |
| Lymphocyte Abs. | Coulter Inst. range | 1.04 - 3.42 | 90% | 1.33 - 3.66 | 100% | Acceptable |
| Monocytes Abs. | Coulter Inst. range | 0.20 - 0.61 | 95% | 0.20 - 0.76 | 95% | Acceptable |
| Eosinophils Abs. | Coulter Inst. range | 0.05 - 1.47 | 95% | 0.05 - 0.73 | 100% | Acceptable |
| Basophils Abs. | Coulter Inst. range | 0.01 - 0.14 | 100% | 0.01 - 0.08 | 95% | Acceptable |
| RBC 1012/L | Tanzanian | 3.64 - 5.93 | 90% | 3.22 - 5.01 | 90% | Acceptable |
| Hemoglobin (g/dL) | Tanzanian | 11.5 - 15.8 | 90% | 9.6 - 14.9 | 95% | Acceptable |
| Hematocrit % | Tanzanian | 33.1 - 45.9 | 95% | 27.4 - 42.7 | 95% | Acceptable |
| MCV (fL) | Coulter Inst. range | 70 - 94 | 95% | 73 - 94 | 95% | Acceptable |
| MCH (pg) | Coulter Inst. range | 23.3 - 33.1 | 95% | 24.5 - 32.9 | 90% | Acceptable |
| MCHC (g/dL) | Coulter Inst. range | 33.1 - 35.6 | 95% | 33.1 - 35.6 | 90% | Acceptable |
| RDW % | Coulter Inst. range | 10.8 - 15.9 | 90% | 11.2 - 17.2 | 100% | Acceptable |
| Platelets (K/uL) | Ugandan | 61 - 345 | 95% | 103 - 383 | 95% | Acceptable |
| MPV (fL) | Coulter Inst. range | 7.2 - 10.2 | 95% | 7.2 - 10.3 | 100% | Acceptable |

Coagulation

Adult:

| Analyte | Adult Reference Ranges |  | % Verified (Expected >90%) |
| --- | --- | --- | --- |
| Pt |  | 11.0-14.0 seconds | 100% |
| Aptt |  | 27-40 seconds | 100% |
| Fibronogen |  | 156-400 mg/dl | 100% |

Pediatric:

| Analyte | Age | Normal Range | % Verified (Expected >90%) |
| --- | --- | --- | --- |
| Pt(sec) | Preterm Infant (30-36 week) | 14.6-16.9 | *** |
|  | Newborn-term | 10.1-15.9 | *** |
|  | 1-5 year | 10.6-11.4 | *** |
|  | 6-10 year | 10.1-12.1 | *** |
|  | 11-16 year | 10.2-12.0 | *** |
| aPtt(sec) | Preterm Infant (30-36 week) | 80-168 | *** |
|  | Newborn-term | 31.3-54.3 | *** |
|  | 1-5 year | 24-36 | *** |
|  | 6-10 year | 26-36 | *** |
|  | 11-16 year | 26-37 | *** |
| Fibrinogen (mg/dl) | Preterm Infant (30-36 week) | 150-373 | *** |
|  | Newborn-term | 167-309 | *** |
|  | 1-5 year | 170-405 | *** |
|  | 6-10 year | 157-400 | *** |
|  | 11-16 year | 154-448 | *** |

***Ranges to be verified over time as appropriate data becomes available.

Reference Range Approval

☐ Approved

☐ Not Approved

| Medical Director: |  |
| --- | --- |
| Date: |  |

| Prepared by: |  |
| --- | --- |
| Date: |  |

Method Approval

☐ Approved

☐ Not Approved (provide recommendations/corrective actions below)

Additional comments, if needed:

| Laboratory Director: |  |
| --- | --- |
| Date: |  |